PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Single channel or dual channel Online Total Organic Carbon Analyser

Analyser for measuring Total Organic Carbon, Total Carbon, Total Inorganic Carbon, Biological Oxygen Demand and/ or Chemical Oxygen Demand via correlation, and Volatile Organic Carbon via calculation. The Analyser is designed for measurement of these parameters in water. The Total Organic Carbon measurement method is infrared measurement of Carbon Dioxide after oxidation. The oxidation method is Two Stage Advanced Oxidation which utilizes the Hydroxyl radical and a manganese catalyst.

The Analyser is capable of measuring Total Organic Carbon every 5 minutes. The Analyser can be configured to measure any combination parameters in up to 2 process streams. The sample sequencer is internal to the Analyser. The Analyser is self-cleaning. The Analyser uses the internal sample pump to bring the sample into the reactor. The Analyser reactor is self-cleaning. The Analyser has internal sensors and utilises self-diagnostics to monitor Analyser status and produces alerts when maintenance is required.

Technical Data

Parameter	TOC, TIC, TC, VOC, after correlation COD, BOD

Measurement method	Infrared measurement of CO2 after oxidation

Oxidation method	Unique Two-Stage Advanced Oxidation Process (TSAO) using Hydroxyl Radicals

Measuring range	0 - 5000 µg/L C

Multi-Stream	Up to 2 process streams and grab sample

Repeatability	± 2 % of reading or ± 10 µg/L C, whichever is greater

Standard Output	4 programmable 4-20 mA analog output signal channel fitted as standard

Digital Output	6 freely programmable system relays

Communication: digital	Modbus, Profibus, Ethernet industrial interface options are available.

Protection class	IP44, standard fan cooled, maximum ambient temperature 45 °C IP54, air cooled, maximum ambient temperature 35 °C

IP54, vortex cooled, maximum ambient temperature 50 °C

EExp / Hazardous

Location

Sample inlet temperature

Certification options are available to European Standards (ATEX for Zone 2) and to North

American Standards (Class I Division 1 and Class I Division 2). Other options are available on request.

2 - 60 °C

Ambient temperature	5 - 45 °C

Cooling and heating options are available.

Humidity	5 - 85 % (non-condensing)

Particle size	Up to 100 µm

Data storage	Previous 9999 analysis data on screen in the microcontroller memory and storage of data archive for the lifetime of the analyser in the SD/MMC card.

Previous 99 fault data on screen in the microcontroller memory and storage of fault data archive for the lifetime of the analyser in the SD/MMC card.

Display	High contrast 40 character x 16 line backlit LCD with LED backlight

User interface	Microcontroller with membrane keyboard

Power requirements

(Voltage)

115 V AC/230 V AC

Power requirements (Hz)   50/60 Hz

Service interval	6 months service intervals

Dimensions (H x W x D)	1000 mm x 500 mm x 320 mm

Weight	50 kg (enclosure weight may change depending on system optional features)

Scope of delivery

[ ] (delete what is not needed)

Analytical Instrument

[ ] Single channel Online Total Organic Carbon Analyser, 0 - 5 mg/L C, [ ] Dual channel Online Total Organic Carbon Analyser, 0 - 5 mg/L C

[ ] Single channel Online Total Organic Carbon Analyser, IP54 purge ready, 0 - 5 mg/L C, [ ] Dual channel Online Total Organic Carbon Analyser, IP54 purge ready, 0 - 5 mg/L C

[ ] Single channel Online Total Organic Carbon Analyser, ATEX compliant, EExp Zone 2, 0 - 5 mg/L C, [ ] Dual channel Online Total Organic Carbon Analyser, ATEX compliant, EExp Zone 2, 0 - 5 mg/L C,

[ ] Single channel Online Total Organic Carbon Analyser, ETL certified, EExp Class 1 Div 2, 0 - 5 mg/L C, [ ] Dual channel Online Total Organic Carbon Analyser, ETL certified, EExp Class 1 Div 2, 0 - 5 mg/L C

Accessories and options

[ ] VOC mode used for measurement of Volatile Organic Carbon in the sample

[ ] Grab sample/calibration port

[ ] Remote activation of manual/calibration port

[ ] Modbus

[ ] Profibus

[ ] Ethernet

[ ] Network control unit

[ ] Sample Sensor to detect presence of the sample and evaluate the quantity of air bubbles in the sample

[ ] Sample overflow chamber

[ ] Flow through sampling chamber with sand trap

[ ] Sampling chamber with sand trap

[ ] Compressor

[ ] Filter pack for removal of water, dirt, and oil from compressed air (if using instrument air supply)

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Hach

Product:	B3500ul Online TOC Analyser