PART 1   GENERAL

1.1	Section includes

A.  Alkalinity process analyzer for continuous monitoring of alkalinity in water.

1.2	Measurement Procedures

A.  The method of measuring alkalinity will be with a colorimetric titration method to determine concentration.

1.3	Alternates

A.  Other methods of alkalinity measurement, such as probes that use a selective membrane or those that use electrolytes, are not acceptable.

1.4	System Description

A.  Performance Requirements

1.   Measurement range

a.	Total alkalinity as calcium carbonate: 1 to 500 mg/L

b.   Phenolphthalein alkalinity as calcium carbonate: 5 to 250 mg/L

2.   Accuracy: ± 5 percent of reading or ±1.0 mg/L, whichever is greater

3.   Repeatability: ±3% of reading or ±0.6 mg/L, whichever is greater

4.   Response time: less than 10 minutes for 90% response to step change at sample inlet (single channel instruments)

5.   Cycle time: 8 minutes (average)

6.   Detection limit: less than or equal to 0.10 mg/L

1.5	Certifications

A.  Safety standards

1.   UL 3101-1

2.   CSA C22.2 No. 1010.1

3.   EN61010-1 (IEC 1010-1)

B.  Class A limits for radio and noise emission as specified by the FCC and EN55011 (CISPR11).

1.6	Environmental Requirements

A.  Operational Criteria

1.   Operating and sample temperature: 5 to 50 °C (41 to 122 °F)

2.   Sample pressure: 0.5 to 30.0 psig (0.03 to 2.04 bar)

3.   Sample flow: 100 to 2000 mL/min. maximum

4.   Sample inputs: up to two sample streams

1.7	Warranty

A.  The analyzer includes a one-year warranty from the date of shipment.

1.8	Maintenance Service

A.  Scheduled maintenance:

1.   Monthly

a.	Reagents and standards:

1)   Replace.

2)   Clean reagent compartment and tubing.

b.   Cleaning solution: fill cleaning solution container (every two weeks in typical municipal wastewater applications).

c.	Autoburette module:

1)   Inspect for seal or fitting leaks.

2)   Inspect for particulate build up.

3)   Lubricate lead screw and ceramic piston guide.

d.   Valve module: inspect module and associated tubing for leaks. e.	Sample conditioning filter:

1)   Inspect filter.

2)   Check sample flow.

3)   Clean or replace the filter.

f.	Mixer module: inspect for particulate build up.

2.   Every three months

a.	Autoburette module: replace piston seals and O-rings. b.   Valve module:

1)   Replace valve rotor

2)   Dry and inspect condition of stator (if scored, replace).

3)   Check for leaks. Replace as needed.

3.   Every six months

a.	Autoburette module:

1)   Check for need to replace piston seals.

2)   Inspect for signs of leakage.

b.   Tubing and fittings: inspect for leaks or damage. Replace as needed. B.  Unscheduled maintenance

1.   Clean instrument enclosure.

2.   Fuse replacement.

PART 2   PRODUCTS

2.1	Manufacturer

A.  Hach Company, Loveland, CO

1.   APA 6000™ Alkalinity Process Analyzer

2.2	Manufactured Unit

A.  The APA 6000 alkalinity analyzer consists of microprocessor-controlled analyzer designed to continuously monitor alkalinity in a sample stream.

2.3	Equipment

A.  The analyzer uses m-cresol purple and bromcresol green indicators for colorimetric measurement of alkalinity at a wavelength of 600 nm.

B.  The analyzer has a digital display in a numeric or graphical format.

C.  The analyzer is capable of automatic calibration, cleaning, and self-priming.

D.  Samples are continuously purged to assure fresh sample to the analyzer and reduce analysis lag time. E.   An automatic burette is used to dispense metered volumes of sample, standards, and reagents.

F.   Sample, standard, and reagent flow are directed to the detector module by a rotary valve.

G.  Grab-sample (10 mL) analysis is possible without interrupting continuous sample flow to the analyzer.

H.  The anayzer is equipped with the following communications capabilities.

1.   Fourteen user-defined internal recorders, of which four can be used for PID control.

2.   Two user-selectable recorder/controller outputs of 4-20 mA, with expansion capability up to 14.

3.   Recorder output span is user-adjustable over the entire span of the analyzer.

4.   Fourteen user-defined alarms. Alarms may be programmed for sample concentration alarms, analyzer system warning, and analyzer system shutdown.

5.   Two unpowered SPDT relays, with expansion capability up to 14, for internal alarms.

6.   Two relay contacts rated for 5 A resistive load at 230 Vac.

I.	Analyzer components are assembled to a NEMA-4X(indoor)/IEC 529 (IP66) plastic enclosure.

J.	All standards and reagents are isolated from the analyzer electronics in separate drip-proof plastic containers.

K.  Power requirement are 95 to 240 Vac, 50/60 Hz.

2.4	Components

A.  Standard equipment:

1.   AquaTrend Interface

2.   One month supply of reagents

3.   Basic sample conditioning kit

4.   Maintenance kit

5.   User manual

B.  Dimensions:  522 x 627 x 526 mm (21 x 25 x 21 inches) C.  Weight:  25.5 kg (56 pounds)

2.5	Accessories

A.  Cable Termination Kit

B.  Digital Display Module (DDM) C.  PS2401 Power Supply

D.  Serial Input/Output Module (SIO) E.   Signal Output Module (SOM)

F.   Installation Kit

G.  Tool Kit

H.  Maintenance Kit

PART 3   EXECUTION

3.1	Preparation

A.  Mounting: bench, wall or panel mount

B.  Sample filter inlet: 3/4-inch NPT male or female

C.  Drain:  gravity, air break, or vent recommended D.  Drain fitting: 3/4-inch NPT barbed hose fitting E.   Data communications distance:

1.   Maximum note-to-node distance: 400 m (1312 feet)

2.   Maximum total wire length: 500 m (1640 feet)

3.2	Installation

A.  Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

B.  Manufacturer’s representative will include a half-day of start-up service by a factory-trained

technician, if requested.

1.   Contractor will schedule a date and time for start-up.

2.   Contractor will require the following people to be present during the start-up procedure. a.	General contractor

b.   Electrical contractor

c.	Hach Company factory trained representative d.   Owner’s personnel

e.	Engineer

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

END OF SECTION