- PRODUCTS

  - Manufacturer

    - Hach Company, Loveland, CO

      - Model DR 6000 Laboratory Spectrophometer

  - Manufactured Unit

    - The power requirements for the DR 6000 spectrophotometer 100/240 Vac, 50/60 Hz.

    - The wavelength capability is 190-1100 nm.

    - The characteristics of the optical system are:

      - Accuracy 5 mAbs at 0.0-0.5 Abs, <1% at 0.5-2.0 Abs at 546 nm

      - Linearity: 0.01 nm at <2 Abs (with neutral glass at 546 nm)

      - Bandwith: 2nm

      - Wavelength accuracy: +/- 1 nm

      - Wavelength Reproducibility: < 0.1 nm

    - Readout modes of the DR 6000 are:

      - Transmittance (%)

      - Absorbance

      - Concentration

      - Scanning wavelength

    - The DR 6000 is equipped with a USB interface, type A (2), USB type B

      - RFID module optional

    - The sample compartment and cell compatibility of the DR 6000 is as follows.

      - 13 mm round

      - 16 mm round

      - 1-cm & 5 cm rectangular

      - 1” round

      - 1” rectangular

      - 10 cm rectangular optional

    - The DR 6000, depending on the test selection, automatically selects the appropriate wavelength.

    - The DR 6000 can store up to 5000 data points (date, time, results, sample ID, user ID) and 200 user- programs.

    - The DR 6000 is capable of measuring the following parameters.

      - Alachlor: 0.1 to 0.5 ppb

      - Alkalinity: 250 to 400 mg/l

      - Aluminum: 0.002 to 0.800 mg/L

      - Ammonia, nitrogen: 0.015 to 50.0 mg/L

      - Arsenic: 0.020 to 0.200 mg/L

      - Atrazine: 0.5 to 3.0 ppb

      - Barium: 2 to 100 mg/L

      - Benzotriazole: 0.2 to 16.0 mg/L

      - Boron: 0.02 to 14.0 mg/L

      - Bromine: 0.05 to 4.50 mg/L

      - Cadmium: 1.3 µg/L to 0.30 mg/L

      - Carbohydrazide: 5 to 600 µg/L

      - Chloramine, mono-: 0.04 to 10.0 mg/L

      - Chloride: 0.1 to 25.0 mg/L

      - Chlorine dioxide: 0.01 to 1000 mg/l

      - Chlorine, free: 0.02 to 10.0 mg/L

      - Chlorine, total: 2 µg/L to 10.0 mg/L

      - Chromium, hexavalent: 0.010 to 1.00 mg/L

      - Chromium, total: 0.01 to 0.70 mg/L

      - Cobalt: 0.01 to 2.00 mg/L

      - Color: 0 to 500 units

      - COD (Chemical Oxygen Demand): 0.7 to 15,000 mg/L

      - Copper: 1 µg/L to 8.00 mg/L

      - Cyanide: 0.002 to 0.240 mg/L

      - Cyanuric Acid: 5 to 50 mg/l

      - DEHA (Diethylhydroxylamine): 3 to 450 µg/L

      - Dissolved oxygen: 6 µg/L to 40 mg/L

      - Erythorbic acid (isoascorbic acid): 13 to 1500 µg/L

      - Fluoride: 0.02 to 2.00 mg/L

      - Formaldehyde: 2 to 500 µg/L

      - Hardness, total (calcium and magnesium as CaCO3): 4 µg/L to 4.00 mg/L

      - Hydrazin: 4 to 600 µg/L

      - Hydroquinone: 9 to 1000 µg/L

      - Iodine: 0.07 to 7.00 mg/L

      - Iron, ferrous: 0.02 to 3.00 mg/L

      - Iron, total: 0.009 to 6.00 mg/L

      - Lead: 3µg/L to 2 mg/l

      - Manganese: 0.006 to 20.0 mg/L

      - Mercury: 0.1 to 2.5 µg/L

      - Methylethylketoxime (MEKO): 15 to 1000 µg/L

      - Molybdenum, molybdate: 0.02 to 40.0 mg/L

      - Nickel: 0.006 to 6 mg/L

      - Nitrate, nitrogen: 0.01 to 35.0 mg/L

      - Nitrite, nitrogen: 0.002 to 250 mg/L

      - Nitrogen Simplfied Total Kjeldahl: 0 to 16 mg/l

      - Nitrogen, total: 0.5 to 150 mg/L

      - Nitrogen, total inorganic: 0.2 to 25.0 mg/L

      - Nitrogen, total Kjeldahl: 1 to 150 mg/L

      - Ozone: 0.01 to 1.50 mg/L

      - PCB (polychlorinated biphenyls): 1 to 50 ppm

      - Phenols: 0.002 to 0.200 mg/L

      - Phosphonates: 0.02 to 125.0 mg/l

      - Phosphorus, acid hydrolyzable: 0.06 to 100 mg/L

      - Phosphorus, reactive (orthophosphate): 19 µg/L to 100 mg/L

      - Phosphorus, total: 0.06 to 100 mg/L

      - Potassium: 0.1 to 7.0 mg/L

      - Quaternary ammonium compounds: 0.2 to 5.0 mg/L

      - Selenium: 0.01 to 1.00 mg/L

      - Silica: 3 µg/L to 100 mg/L

      - Silver: 0.005 to 0.700 mg/L

      - Sulfate: 2 to 900 mg/L

      - Sulfide: 5 to 800 µg/L

      - Surfactants, anionic: 0.002 to 0.275 mg/L

      - Suspended solids: 5 to 750 mg/L

      - Tannin and lignin: 0.10 to 0.90 mg/L

      - TOC (Total Organic Carbon): 0.3 to 700 mg/L

      - Tolyltriazole: 1.0 to 20.0 mg/L

      - Toxicity: 0 to 100% Inhibition

      - TTHM (trihalomethanes, total): 10 to 600 µg/L

      - TPH (Total Petroleum Hydrocarbons):  2 to 200 ppm,threshold

      - Volatile Acids: 27 to 2800 mg/L

      - Zinc: 0.01 to 3.00 mg/L

    - Manufacturer shall provide high quality pre-mixed reagents and standards.

  - Instrument includes

    - Universal Cell Adapater

    - Manual

    - Power supply

    - Dust Cover

    - 1 inch square matched glass sample cells

  - Accessories

    - USB Barcode Scanner (if specified)

    - USB Keyboard (if specified)

    - Test Filter set (if specified)

    - Sipper units for flow thru application (if specified)

    - Multi-cell holder for 5x1 Inch square cells (if specified)

    - Multi-cell holder for 7x1 cm square cells (if specified)

  - Dimensions

    - Length: 460 mm (18.11 in)

    - Width: 500 mm (19.69 in)

    - Height: 215 mm (8.46 in)

    - Weight: 11 kg (24.25 lb)

- EXECUTION

  - Preparation

    - None required.

  - Installation

    - None required.

- 2.3       Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

- END OF SECTION