PART 1   GENERAL

1.1	Section includes

    - Analyzer for monitoring of phosphate in water.

    - Includes the capability to remotely monitor sensors on any browser-enabled device and present

diagnostics on the overall health of the measurements (on Predictive Diagnostics-enabled sensors),

as well as upcoming and required maintenance - reducing user risk and downtime. It includes the

capability to connect to a laboratory spectrophotometer to correct process measurements based on lab

samples, without having to remove the process sensor from the water.

1.2	Measurement Procedures

A.  The method of measuring ammonia will be by vanadomolybdophosphoric acid colorimetric method as found in Standard Methods.

B.  At the beginning of each measurement cycle the analyzer measures the sample for purposes of compensating for background color of the sample.

1.3	Alternates

- Other instruments that do not use the vanadomolybdophosphoric acid colorimetric method are not acceptable and that do not compensate for background sample color are not acceptable

- Other instruments that do not have predictive diagnostic capabilities are unacceptable

1.4	System Description

A.  Performance Requirements

      - Measurement range (depending on model):

        - 0.05 to 15 mg/L, or b.   1 to 50 mg/L

      - Lower detection limit (depending on model):

        - 0.05 mg/L, or b.   1 mg/L

      - Accuracy (depending on model):

        - 2% ±0.05 mg/L, or b.   2% ±1 mg/L

      - Reproducibility (depending on model):

        - 2% ±0.05 mg/L, or b.   2% ±1 mg/L

      - Response time: Less than 5 minutes (T90), including sample preparation

      - Measurement interval: 5 to 120 minutes, adjustable

      - When connected to a multi-parameter digital control the overall status of the instrument performance is displayed as a percentage value via a measurement indicator

      - When connected to a multi-parameter digital control the overall time remaining until maintenance tasks are due is displayed in days

1.5	Certifications

- none

1.6	Environmental Requirements

A.  Operational Criteria

1.   Sample temperature: 4 to 40 °C (39 to 104 °F)

2.   Sample pH: 5 to 9 pH

3.   Operating temperature: -20 to 45 °C (-4 to 114 °F)

4.   Operating humidity: 95% relative humidity, non condensing

1.7	Warranty

A.  The products include a one-year warranty from date of shipment.

1.8	Maintenance Service

A.  Unscheduled maintenance

1.   Replace chemicals as required

PART 2   PRODUCTS

2.1	Manufacturer

A.  Hach Company, Loveland, CO

1.   Model PHOSPHAX sc Phosphate Analyzer

2.2	Manufactured Unit

A.  The PHOSPHAX sc Phosphate Analyzer consists of:

1.   Housing: ASA UV-resistant, IP55-rated, lockable

2.   Gas sensitive electrode

3.   Colorimeter

2.3	Equipment

A.  The analyzer calibrates and cleans itself automatically.

B.  The required power supply is 230 Vac/50Hz or optional 115 Vac/50-60Hz connected to a Hach model  sc1000 multi-parameter universal controller.

C.  Data transmission is made with a data cable with the controller.

D.  Outputs include relay, current outputs, and bus interface via the controller.

2.4	Components

A.  Standard equipment:

1.   Analyzer

2.   Manual

3.   Reagents

B.  Dimensions: 21.3 x 28.3 x 15.4 inches (540 x 720 x 390 mm) C.  Weight: 77 pounds (35 kg)

2.5	Accessories

A.  Hach Filterprobe sc

B.  Hach sc1000 multi-parameter universal controller

C.  Mounting kits (rail or stand)

PART 3   EXECUTION

3.1	Preparation

A.  Wall-mount indoors or outdoors. Rail- and stand-mounting options available.

3.2	Installation

A.  Contractor will install the sensor in strict accordance with the manufacturer’s instructions and recommendation.

B.  Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

1.   Contractor will schedule a date and time for start-up.

2.   Contractor will require the following people to be present during the start-up procedure. a.	General contractor

b.   Electrical contractor

c.	Hach Company factory trained representative d.   Owner’s personnel

e.	Engineer

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

END OF SECTION