PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Online Streaming Current Monitor

Online streaming current monitor sensor for measurement of charge on particles in the coagulation phase of water, wastewater, or industrial treatment processes. Sensor construction is stainless steel, aluminum, epoxy, and UHMW. Sensor is capable of measuring as frequently as one second and is equipped with self-diagnostic routines. Communication via 4-20mA outputs and alarm relays. Sensor can be equipped with auto-flush mechanism to decrease the need for manual cleaning.

Technical Data

Compliance

certifications

CE, cETLus (for UL 61010-1 &C22.2. No. 61010-1 safety standards)

Response time	1 sec to 1 min adjustable average

Flow rate	Variable; 2 - 4 L/min recommended

Humidity	Maximum relative humidity 95% for temperatures up to 50 °C (122 °F)

Sample temperature	7 - 35 °C (45 - 95 °F)

Operating temperature range

0 - 50 °C ; 95% relative humidity (non-condensing)

Storage conditions	-20 °C - 70 °C ; 95% relative humidity (non-condensing)

Connections	1/2 inch BSP

Inputs	Dual digital (2x); optically isolated; switch input to 5 - 24 V Output	4-20 mA; Maximium load 500 Ω, eletrically isolated Relays	Relay (2x); SPST; 5 A, 30 V DC resistive loads Communication	Optional: 2x 4-20 mA Outputs

Pollution degree	2, II

Protection class	I

Power requirements

(Voltage)

110 V AC

Wetted materials	Stainless Steel, UHMW, Epoxy Dimensions (H x W x D)	750 mm x 320 mm x 250 mm Weight	21.5 kg

Scope of delivery

[ ] (delete what is not needed) Analytical Instrument

[ ] Online Streaming Current Monitor

Accessories

[ ] Water connections kit with grit filter

[ ] Auto-flush Mechanism with Auto-flush communications card

[ ] 4-20 mA Output Card

[ ] Relay Output Card

[ ] 4-20 mA PID Card

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Hach

Product:	AF7000 Streaming Current Monitor