Nitrate Sensor with predictive diagnostics

Sensor for continuous measuring of NO3-N/NO3 in water. The method of measuring is potentiometric by ion-selective electrodes (ISE). A differential pH electrode is used as the reference. A chloride ISE is used to correct the nitrate value for the presence of chloride ion. The sensor requires a universal digital controller. The NO3 sensor consists of a digital sensor with integrated, replaceable sensor cartridge. The integrated sensor cartridge includes nitrate ISE, differential pH electrode, chloride ISE, and temperature sensor. The sensor cartridge is factory calibrated with all electrodes individually calibrated and also calibrated to each other. The sensor performs one- or two-point inline matrix corrections to adapt it to the medium matrix. The sensor is fully immersible and does not require sample conditioning.

The sensor has built in predictive diagnostics capability which increases available operation time. The predictive diagnostics produces alerts for upcoming maintenance tasks by monitoring the instrument’s internal components and tracking service requirements. When connected to a properly equipped controller the controller will display status indicators for each sensor with predictive diagnostics capability.

Capable of performing matrix correction using reference data generated in and transferred from a laboratory spectrophotometer. The process probe allows for a secure two way communication. It can send data to be stored with the photometric reference data or receive information to perform a matrix correction of the process probe by using the data generated in the laboratory.

Includes the capability to remotely monitor sensors on any browser-enabled device and present diagnostics on the overall health of the measurements (on predictive diagnostics-enabled sensors), as well as upcoming and required maintenance - reducing user risk and downtime. It includes the capability to connect to a laboratory spectrophotometer to correct process measurements based on lab samples, without having to remove the process sensor from the water.

Technical Data

Measurement method

Measuring range pH  range Accuracy Response time Calibration method

Potentiometric ion-selective measurement 0 - 1000 mg/L NO3-N , Cl-

5 - 9 pH

±5 % of measured value + 0.2 mg/L (with standard solution) NO3-N

< 3 min

With Cartrical plus technology: automatic import of factory calibration data from cartridge to probe by RFID; 1 and 2 point matrix correction

2 - 40 °C

Air: -20 - 45 °C

Water: 2 - 40 °C

With Cartrical plus technology: compact housing containing calibrated electrodes for nitrate and chloride, reference system and temperature sensor, all calibrated to each other; typical lifetime 12 months

< 4 m/s max.

Probe: Stainless Steel 1.4571; ASA + PC, silicon; PVC and PU;

Sensor cartridge: PVC; POM; ABS; Stainless Steel 1.4571; NBR; Continuous

0.2 mg/L NO3-N

84.5 mm x 320 mm

45 ° ±15 ° (vertical in flow direction)

10 m fixed cable at sensor, extendable to 100 m

Universal Controller, Multi-Parameter Universal Controller

Sample temperature

Operating temperature

range

Sensor cartridge

Flow

Wetted materials

Measuring interval Detection limit Dimensions (D x L) Installation angle Cable length

Controller

compatibility Process connection Protection class Sample pressure Storage conditions

Weight

1 inch thread

IP 68

max. 0.3 bar

5 - 40 °C sensor cartridge

2.4 kg

PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Scope of delivery

[ ] (delete what is not needed)

Analytical Instrument

[ ] Nitrate Sensor with RFID (Not available in Cyprus, Macedonia, Australia, New Zealand, Russia, Ukraine, Africa, Asia, America)

[ ] Nitrate Sensor without RFID

Accessories

[ ] Polishing paper for chloride electrode

[ ] Digital extension cable, with molded plug and coupling	(select 5/10/15/20/30/50 m) [ ] Cleaning unit for sensor without compressor

[ ] Solenoid valve with pressure gauge (used for relay-controlled air pressure cleaning) [ ] Air blast cleaning system cable kit- including power- and relay cable

[ ] Air blast cleaning system

Mounting Hardware

[ ] Mounting hardware as follows:

[ ] Rim mounting kit for sensor, made of stainless steel [ ] Rail mounting kit for sensor, made of PVC

[ ] Chain Mounting for sensor, made of PVC [ ] Sensor adapter for rail mounting unit

[ ] Additional weight for cleaning unit for use with cleaning unit and chain mounting kit [ ] Adapter 1½” NPT - 1” NPT

[ ] Fitting, 45° angle, 2 x 1½” NPT connectors, made of PVC

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:

Hach

Product:

N-ISE sc Nitrate Sensor with Prognosys predictive diagnostics