PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

High Purity Water pH & ORP Panel

High purity water pH sensor specially designed for continuous in-stream pH measurement of pure and ultra-pure water. It uses a reference electrode very similar to the measuring electrode (combined concentric electrode) to minimise the impedance under which the measurement is carried out.

The system is completely sealed to prevent contamination from atmospheric carbon dioxide. Several design features are used to supply online pH/ORP measurements with a high degree of accuracy. The panel connects to a digital controller for operation.

Technical Data

Dimensions (H x W x D)	305 mm x 384 mm x 166 mm

Weight	3.6 kg

Flow rate	100 - 300 mL/min Operatin conditions	4 bar maximum at 25 °C Temperature sensor	Pt100

Temperature range	0 - 80 °C

Operating temperature range	0 - 60 °C

Operating humidity	0 - 100 % (non-condensing)

Measuring range	2 - 12 pH at 0 - 80°C

-1500 - 1500 mV at 0 - 50 °C Accuracy	<0.1 pH or ±5 mV for ultra pure water Repeatability	0.01 pH at 25 °C (constant flow) Cable length	7.7 m, supplied

Certifications	CE

Scope of delivery

[ ] (delete what is not needed)

The panel is assembled at the factory with all the components attached and plumbed.

Accessories

[ ] 4 pH solution, NIST, 500 mL [ ] 7 pH solution, NIST, 500 mL

[ ] 10 pH solution, NIST, 500 mL

[ ] ORP reference solution, 200 mV, 500 mL [ ] ORP reference solution, 600 mV, 500 mL [ ] Controller extension cable, 1 m

[ ] Controller extension cable, 7.6 m

[ ] Controller extension cable, 15.2 m

[ ] Controller extension cable, 30.5 m

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Polymetron

Product:	8362 High Purity Water pH & ORP Panel