PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Total Free Chlorine Amperometric Sensor

On-line, single-channel industrial sensor that measures free chlorine in drinking water treatment plants, distribution networks, and other applications that require monitoring free chlorine at the ppb and ppm levels. The instrument uses an amperometric method to measure HOCl concentration and has a low minimum detection limit and a wide measuring range. A membrane allows the selective diffusion of HOCl molecules to the amperometric sensor. The sensor requires a universal digital controller.

Technical Data

Measuring range	0 - 20 mg/L HOCl

Detection limit	0.005 mg/L HOCl

Accuracy	2 % or ±10 ppb HOCl, whichever is greater at pH <8

Standard deviation	0.7 %

Response time	90 % in less than 90 seconds

Measuring interval	Continuous

Flow	14 L/h (200 - 250 mL/min) auto-regulated by flow through cell

Pressure range	0.1 - 2 bar in flow cell

Sample temperature	2 - 45 °C

Temperature compensation

Automatic over sample temperature range

pH Range	4 - 8 pH, acidification unit available for pH >8

Measuring principle	Amperometric/Membrane (electrode, membrane, electrolyte)

Interferences	No interference from chloramines.

Chlorine dioxide and ozone interfere with measurement.

Calibration method	Zero calibration: electrically (automatic) or with de-chlorinated water

Calibration: comparison of lab method with process sample

pH calibration: single- or two-point using standards or comparison of lab method with process sample

Calibration interval	2 months

Mounting	Flat, vertical surface (panel, stand, etc.)

Sample input	0.25” outer diameter Sample output	0.5” inner diameter Storage conditions	-20 °C - 60 °C

Relative humidity	0 - 90 % non-condensing

Maintenance interval   Measurement cell: 6 months for membrane and electrolyte for typical application pH: 1 - 1.5 years for typical application

Height	270 mm

Width	250 mm

Weight	6.5 kg

Enclosure waterproof rating

Controller

compatibility

IP66 / NEMA 4X

Universal controller, Multi-parameter universal controller

Scope of delivery

[ ] (delete what is not needed) Analytical Instrument

Total Free Chlorine Amperometric Sensor

Optional Accessories

[ ] Acidification unit

[ ] Intermittent flow unit

[ ] Instrument stand

[ ] 125 V Power cord with strain relief

[ ] 230 V Power cord with strain relief

[ ] Digital extension cable  	(select 1/7.6/15.2/30.5 m) (Standard cable length is 0.4 m)

[ ] Digital termination box (required when the length of cable between the digital sensor and controller is between 100 and 1000 m)

Replacement Parts

[ ] pH electrode

[ ] Chlorine electrode

[ ] Pre-mounted membranes (pk/4) [ ] Electrolyte

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Hach

Product:	9184 sc Total Free Chlorine Amperometric Sensor