PART 1	GENERAL

1.1	Section includes

- Sensor that continuously measures pH in aqueous solutions.

- Includes the capability to remotely monitor sensors on any browser-enabled device and present diagnostics on the overall health of the measurements (on Predictive Diagnostics-enabled sensors), as well as upcoming and

required maintenance - reducing user risk and downtime.

1.2	Measurement Procedures

1.	The method of measuring pH will be with a probe or sensor that uses differential electrode measurement technique using three electrodes.

a.	Two electrodes compare the process value to a stable internal reference standard buffer solution. b.	The internal electrode is non-flowing, foul-resistant characteristics.

1.3	Alternates

A.   Other methods of pH measurement including combination probes will not be accepted.

B.   Probes or sensors that do not communicate with the Hach SC platform controller will not be accepted.

1.4	System Description

A.   Performance Requirements

1.	Measurement range: 0 to 14 pH

2.	Sensitivity: 0.01 pH

3.	Stability: 0.03 pH per 24 hours, non-cumulative

1.5	Certifications

A.   General Purpose CSA/CSANRTL and FM (UL Pending) when used with an approved controller.

B.   Class 1, Div 2 Groups A thru D CSA/CSANTRL and FM (UL Pending) when used with an approved controller.

1.6	Environmental Requirements

A.   Operating Criteria

1.	Temperature range: 23 to 158 °F (-5 to 70 °C)

2.	Sample flow rate: 3 meters (10 feet) per second, maximum

3.	Pressure: 100 pounds per square inch at 158 °F (6.9 bar at 70 °C)

4.	Transmission distance: 1000 meters (3240 feet), maximum

1.7	Warranty

A.   The product includes a 30-month pro-rated warranty from the date of shipment.

1.	0 to 12 months: 100% free from manufacturer defects

2.	0 to 18 months: replaced at 1/3 current list price for any reason

3.	19 to 30 months: replaced at 2/3 current list price for any reason

1.8	Maintenance Service

A.   Scheduled maintenance:

1.	Condition/site dependent includes:

a.	Occasional wipe of electrode with mild soap solution b.	Soak sensor for 2 to 3 minutes

c.	Use small bristle brush to scrub measuring end of sensor

2.	Twice yearly: replace salt bridge and pH buffer

PART 2   PRODUCTS

2.1	Manufacturer

A.   Hach Company, Loveland, CO

1.	Hach model pHD-SC pH Sensor

2.2	Manufactured Unit

A.   The Hach pHD-SC pH Sensor consists of:

1.	Encapsulated sensor with digital signal processor, pre-amplifier, reference electrode, measuring electrode, ground electrode, and replaceable salt bridge.

2.	Integral 10-meter cable

3.	Wetted material [select one]:

a.	PEEK®

b.	Ryton®.

2.3	Equipment

A.   The Hach pHD-SC sensor works with Hach SC platform controllers only. B.   The probe has a built in Pt 1000 RTD temperature compensator.

C.   Mounting styles are as follows:

1.	1-inch NPT threads on both ends to mount into:

a.	Standard 1-inch pipe tee

b.	GLI adapter pipe for union mount with standard 1.5-inch tee c.	End of pipe for immersion into vessel

2.	For PEEK® housing only [select one]:

a.	Insertion body style: 1-inch NPT threads only on the cable end to mount into a GLI ball valve hardware assembly so that the sensor can be inserted into or retracted from the process without stop to the process flow.

b.	Sanitary body style: 2-inch flange to mount into GLI 2-inch sanitary tee with special cap and EDPM

compound gasket.

2.4	Components

A.   Standard equipment:

1.	Probe

2.	Salt bridge

3.	10-meter integral cable

4.	Manual

B.   Dimensions

1.	Length: 10.75 in. (28.8 cm)

2.	Diameter 1.25 in. (3.1 cm) C.   Weight: 11 oz. (0.316 kg)

2.5	Accessories

A.   Air blast cleaning system

B.   Junction box for extension cables. Must be used for lengths greater than 100 meters. C.   Extension cables

D.   ModBUS® RS-232 or RS-485 digital output card

E.   Mounting hardware

F.   pH buffer solutions

PART 3	EXECUTION

3.1	Preparation

A.   The sensor must be mounted to a Hach mounting assembly directly in the solution to be measured.

B.   Mount sensor vertically with electrode pointing down. A minimum of 15 degrees above horizontal is allowed.

3.2	Installation

A.   Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

B.   Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

1.	Contractor will schedule a date and time for start-up.

2.	Contractor will require the following people to be present during the start-up procedure. a.	General contractor

b.	Electrical contractor

c.	Hach Company factory trained representative d.	Owner’s personnel

e.	Engineer

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

END OF SECTION