- GENERAL

  - Section includes:

    - Sensor that continuously measures pH or Oxidation/Reduction Potential in aqueous solutions.

  - Measurement Procedures

    - The method of measurement will be determination of the electron activity of a solution by using an inert indicator electrode and a reference electrode. The potential difference between the indicator electrode and the reference electrode equals the redox potential of the system.

  - Alternates

    - Probes or sensors that do not communicate with Hach model sc200 or Polymetron 9500 controllers will not be acceptable.

  - System Description

    - Performance Requirements for 8362 ORP Sensor

      - Measurement Range: ±1500 mV

      - Temperature compensation for sample temperature between -20 to 200 °C (-4 to 392 °F)

  - Certifications

    - EMC: CE compliant for conducted and radiated emissions CISPR 11 (Class A limits), EMC Immunity EN 61326-1 (Industrial limits) when part of an approved system

    - Safety: General Purpose UL/CSA 61010-1 with cETLus safety mark when part of an approved system

    - Australian C-TICK and Korean KC Markings when part of an approved system

  - Environmental Requirements

    - Operational Criteria

      - Operating Temperature: -20 to 60 °C (-4 to 140 °F)

      - Maximum Pressure: 16 bar at 25 °C, 6 bar at 100 °C (232 psi at 77 °F, 87 psi at 212 °F)

      - Relative humidity: 10 to 95%, non-condensing

  - Maintenance Service

    - Scheduled Maintenance

      - Monthly

        - Calibrate Sensor

      - Each 6 months

        - Evaluate Sensor for Replacement

    - Unscheduled Maintenance

      - Clean sensor with appropriate cleaning solution at appropriate intervals depending on the sample tested.

- PRODUCTS

  - Manufacturer

    - Hach Lange Sárl, Geneva, Switzerland

  - Manufactured Unit

    - The 8362 ORP sensor consists of:

      - An ORP sensor composed of stainless steel

      - Integral Cable

      - The sensor is individually factory-tested to determine its individual slope offset.

  - Equipment

    - The Polymetron 8362 ORP Sensor works with Hach model sc200 or Polymetron 9500 controllers only.

    - The probe has a separate Pt 100 temperature sensor.

    - Wetted materials as follows:

      - 8362 Sensor

        - Stainless Steel

        - Guarded Glass

  - Components

    - Standard equipment

      - Sensor

      - Integral Cable

    - Dimensions:

      - 8362 Sensor: 22 mm x 95 mm (0.87in x3.74 in)

    - Weight: 0.3kg (0.66 lbs)

  - Required Accessories

    - Flow-through chamber

    - Temperature sensor

  - Optional Accessories

    - Cables

      - 5m (16 ft)

      - 10 m (33 ft)

      - 20 m (66 ft)

- EXECUTION

  - Preparation

    - The sensor must be connected to a Hach or Polymetron flow-through chamber that is plumbed appropriately to the sample.

  - Installation

    - Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require the following people to be present during the start-up procedure.

        - General contractor

        - Electrical contractor

        - Hach Company factory trained representative

        - Owner’s personnel

        - Engineer

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

- END OF SECTION