PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Digital Inductive Conductivity Sensor

Electrodeless inductive conductivity sensor to continuously measure conductivity and/or resistivity in aqueous solutions. The sensor measures up to 2,000,000 µS/cm and has a built-in Pt 1000 RTD thermistor to compensate for changes in process temperature. Sensor materials of construction can withstand high pressures and temperatures.

The sensor connects to a universal digital controller through a digital gateway.

# Technical Data

| Measuring range | 200 µS/cm - 2000 mS/cm |
| --- | --- |
| Linearity | ≥ 1.5 mS/cm: ±1% of reading < 1.5 mS/cm: ±15 µS/cm |
| Response time | (T ) 0.5 s 90 |
| Accuracy | > 500 µS/cm: ±0.5% of reading < 500 µS/cm: ±5 µS/cm |
| Cable length | 200 - 2000 µS/cm: max. 61 m 2000 µS/cm - 2000 mS/cm: max. 91 m |
| Temperature sensor | Pt1000 RTD |
| Sensor cable | 5 conductor (plus 1 shield), 6 m, rated at 150 °C - polypropylene and PVDF sensors at 200 °C - PEEK and PFA sensors |
| Material | Polypropylene, PFDF, PEEK, or PFA |
| Temperature / Pressure limit | Polypropylene: 100 °C at 6.9 bar PVDF: 120 °C at 6.9 bar PEEK and PFA: 200 °C at 13.8 bar |
| Flow | max. 3 m/s |

Scope of delivery

Sensor complete with cable, digital gateway, extension cable, and manual

Sensor Housing Options [ ] Select at least 1 option

[ ] Polypropylene (Textiles, Natural Waters, Seawater, Steam Generation) [ ] PVDF (Acidic applications, Wastewater)

[ ] PEEK (Pulp/Paper, Wastewater)

[ ] PFA Teflon (Food and Beverage, Chemicals and Refining)

Sensor Mounting Options [ ] Select at least 1 option [ ] Convertible

[ ] Immersion [ ] Union

[ ] Insertion

[ ] Sanitary (not available in PEEK)

Accessories

[ ] (delete what is not needed)

[ ] Conductivity solution, 100-1000 µS/cm, 1 L

[ ] Conductivity solution, 2000-150,000 µS/cm, 1 L [ ] Conductivity solution, 1000-2000 µS/cm, 1 L

[ ] Conductivity solution, 200,000-300,000 µS/cm, 1 L

Services

[ ] (delete what is not needed)

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Hach

Product:	3700 sc Digital Inductive Conductivity Sensor