PLEASE NOTE: The following specification contains areas, highlighted in yellow and with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

- GENERAL

  - Section includes:

    - Instrument for discontinuous, online monitoring of Chloride in water using ISE (ion-selective electrode) technology.

    - EZ3005sc Chloride Analyzer

  - Measurement Procedures

- The analyzer is used with an integrated Controller to measure Chloride in water. The analytical method is direct ISE measurement by combined ion-selective electrode, conform with standard methods EPA 9212 and ASTM D512-12.

  - Alternates

    - Analyzers without standard automatic procedures for calibration, validation and cleaning are not acceptable.

    - Analyzers without the ability for single sample ("grab sample") measurement are not acceptable.

    - Analyzers without option for up to 8 sample streams are not acceptable.

    - Analyzers without analog and digital output options are not acceptable.

    - Analyzers without an integrated industrial panel PC are not acceptable.

  - System Description

    - Performance Requirements

- Measuring Range

- [   ] 2 to 100 mg/L (parts per million)
[   ] 5 to 250 mg/L
[   ] 10 to 500 mg/L
[   ] 20 to 1,000 mg/L

    - Other Specifications

      - Limit of detection

- 2 - 100 mg/L Cl-: 2 mg/L

- 5 - 250 mg/L Cl-: 5 mg/L

- 10 - 500 mg/L Cl-: 10 mg/L

- 20 - 1000 mg/L Cl-: : 20 mg/L

      - Precision / Repeatability

        - Better than 2% full scale range for standard test solutions

      - Cycle Time

        - 5 minutes

      - Cleaning

        - Automatic; Frequency freely programmable: 6hours, 12 hours, daily, weekly

      - Calibration

        - Automatic; 2-point; offset or slope; Frequency freely programmable: 6hours, 12 hours, daily, weekly. Note: manufacturer recommends that a calibration is done when the reagents are replaced

      - Validation

        - Automatic; Frequency freely programmable: 6hours, 12 hours, daily, weekly

  - Certifications

    - CE compliant

    - UL certified

  - Environmental Requirements

    - Operational Criteria

      - Operating temperature: 10 to 30 °C ±4 °C deviation (50 to 86 °F ±7.2 °F deviation)

      - Reagent temperature: keep between 10 to 30 °C ±4 °C deviation(50 to 86 °F)

      - Relative humidity: 5 to 95 %, non-condensing

  - Warranty

    - Warranted from manufacturer defects for two years (Europe) or one year (all other geographies) from date of shipment.

  - Maintenance and Service

    - Unscheduled Maintenance

      - Check and clean the analyzer components, depending on cleanliness of the sample

    - Scheduled Maintenance / preventative

      - Monthly

        - Reagents refill; validation and / or calibration

      - Quarterly

        - Pump tubing and pinch-valve tubing replacement

      - Half-yearly

        - Electrode replacement

      - Annually

        - Replacement of valves and pistons

PRODUCTS

  - Manufacturer

    - Hach

    - EZ3005sc Chloride Analyzer

  - Manufactured Unit

    - The Chloride analyzer consists of a microprocessor-controlled Ion-selective analyzer designed to monitor Chloride discontinuously in a sample stream. Automatic cleaning, calibration and validation are available.

  - Equipment

    - Online Analyzer

      - Utilizes direct measurement by combined ion-selective electrode.

      - With automatic calibration, validation, priming and cleaning

      - If chosen, the instrument can monitor up to 8 sample streams.

      - Plug-and-play airtight fittings for connecting reagent tubing

      - High precision micropumps for reagent dosing

    - Controller

      - Industrial integrated controller.

      - User interface with different user levels

      - If chosen, includes capability to communicate measurements via 4-20 mA outputs, Modbus TCP/IP, Modbus RS485, Profibus, Ethernet IP

    - Reagents and Standards

      - The analyzer shall use quick connect reagent containers with pre-installed tubing.

      - Reagents and standards shall be prepared according to instructions on the Method + Reagent sheet.

  - Components

    - Analytical instrument
To deliver:

      - Chloride Analyzer as selected in section 1.1.A. : will be delivered as standard range and will require range set-up upon commissioning

      - Wall-mount bracket

      - Reagent containers

      - User Manual

      - Method + Reagent Sheet

    - Dimensions: refer to analyzer drawings

    - Weight: 35 kg (77 lb)

  - Instrument Options,

- *******************************************************

- Must be added to instrument at time of order.

- *******************************************************

- Outputs (select / fill in one)

- [     ]4 x 4-20 mA Outputs

- [     ]8 x 4-20 mA Outputs

- [     ]4 x 4-20 mA Outputs + Modbus RTU

- [     ]8 x 4-20 mA Outputs + Modbus RTU

- [     ]4 x 4-20 mA Outputs + Modbus TCP/IP

- [     ]8 x 4-20 mA Outputs + Modbus TCP/IP

- [     ]4 x 4-20 mA Outputs + Profinet

- [     ]8 x 4-20 mA Outputs + Profinet

- [     ]4 x 4-20 mA Outputs + Profibus DP

- [     ]8 x 4-20 mA Outputs + Profibus DP

- [     ]4 x 4-20 mA Outputs + Ethernet/IP

- [     ]8 x 4-20 mA Outputs + Ethernet/IP

  - Instrument Accessories

- ***************************

- Select as many as required

- Number of sample streams (fill in, select up to 8):

- [     ] sample streams

- This to order the moduplex

- APPAZ0080002 - Moduplex, 2 streams, Pinch Valve, 1/8" OD
APPAZ0080004 - Moduplex, 4 streams, Pinch Valve, 1/8" OD
APPAZ0080008 - Moduplex, 8 streams, Pinch Valve, 1/8" OD

- [   ] EZ9010 EZ-Size standard version (immersion version), pore size 50 µm

- [   ] EZ9010 EZ-Size standard version (immersion version), pore size 100 µm

- [   ] EZ9020 EZ-Size standard version (immersion version), pore size 50 µm

- [   ] EZ9020 EZ-Size standard version (immersion version), pore size 100 µm

- [   ] EZ9150 Heavy-Duty Filtration System, pore size 10 µm

- [   ] EZ9150 Heavy-Duty Filtration System, pore size 50 µm

- [   ] EZ9150 Heavy-Duty Filtration System, pore size 100 µm

- [   ] EZ9200 Microsize standard version (immersion version), pore size 0.04 µm

- [   ] EZ9250 Microsize standard version (immersion version), pore size 0.04 µm

- [   ] APPAR0000160 Table Stand

- [   ] EZ950x Reagent Rack

EXECUTION

  - Preparation

      - Mounting

        - As shown on the drawings

      - Inlet and outlet connection sizes

        - As shown on the drawings

      - Sample Flow Rate

        - 100 to 300 mL/minute

      - Sample Pressure

        - By external overflow vessel

      - Sample Temperature

        - 10 to 30 °C(50 to 86 °F)

      - Other sample requirements

        - Maximum particle size 100 µm, <0.1 g/L suspended solids, turbidity <50 NTU

  - Installation

    - Install analyzer following transmittal drawings and instrument user manual.

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel. Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts is required. Third-party parts reagents are not approved for use.

DOC353.53.30545.Dec19

END OF SECTION