**PLEASE NOTE: The following specification contains areas, highlighted in yellow and with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.**

GENERAL

* 1. Section includes:
     1. Instrument for semi-continuous, online monitoring of Alkalinity in water using titration. The determination of Alkalinity shall be for:

[ ] Total Alkalinity and Free Alkalinity

[ ] Free Alkalinity and Total Hardness

[ ] Total Alkalinity and Total Hardness

[ ] Total Alkalinity and Free Alkalinity and Total Hardness

[ ] Total Alkalinity and Free Alkalinity and Total Hardness and Calcium Hardness

* 1. Measurement Procedures

The analytical method to determine Alkalinity in water is titration: For the analysis of Alkalinity only it is an acid-base titration with hydrochloric acid. For the analysis of Hardness this is a colorimetric calmagite EDTA titration with LED dipping probe.

* 1. Alternates
     1. Analyzers without standard automatic procedures for validation and cleaning are not acceptable.
     2. Analyzers without option to measure up to 4 parameters in parallel are not acceptable.
     3. Analyzers without the ability for single sample ("grab sample") measurement are not acceptable.
     4. Analyzers without option for up to 8 sample streams are not acceptable.
     5. Analyzers without analog and digital output options are not acceptable.
     6. Analyzers without an integrated industrial panel PC are not acceptable.
  2. System Description
     1. Performance Requirements
        1. Measuring Range
           1. Total Alkalinity and Free Alkalinity  
              [ ] 10 to 500 mg/L CaCO3   
              [ ] 25 to 1250 mg/L CaCO3   
              [ ] 50 to 2500 mg/L CaCO3   
              [ ] 100 to 5000 mg/L CaCO3
           2. Total Alkalinity and Total Hardness  
              [ ] TA: 10 to 250 mg/L | TH: 10 to 250 mg/L CaCO3   
              [ ] TA: 20 to 500 mg/L | TH: 25 to 500 mg/L CaCO3  
              [ ] TA: 40 to 1000 mg/L | TH: 25 to 1000 mg/L CaCO3
           3. Alkalinity and Hardness  
              [ ] TA/FA: 10 to 500 mg/L | TH/CaH: 10 to 100 mg/L CaCO3   
              [ ] TA/FA: 25 to 1250 mg/L | TH/CaH: 10 to 250 mg/L CaCO3   
              [ ] TA/FA: 50 to 2500 mg/L | TH/CaH: 25 to 500 mg/L CaCO3   
              [ ] TA/FA: 100 to 5000 mg/L | TH/CaH: 25 to 1000 mg/L CaCO3
        2. Limit of detection
           1. Alkalinity: ≤ 10 mg/L
           2. Hardness: ≤ 10 mg/L
        3. Precision
           1. Better than 2% full scale range for standard test solutions
     2. Other Specifications
        1. Cycle Time
           1. 20 – 30 minutes
           2. 40 minutes (+ Calcium Hardness)
        2. Cleaning
           1. Automatic; frequency freely programmable
        3. Calibration
           1. NA
        4. Validation
           1. Automatic; frequency freely programmable
        5. Alarm
           1. 1x malfunctioning, 4x user-configurable, max. 24 VDC/0.5 A, potential free contacts
        6. Protection class
           1. Analyzer cabinet: IP55
           2. Panel PC: IP65
        7. Material
           1. Hinged part: Thermoform ABS, door: plexiglass
           2. Wall section: Galvanized steel, powder coated
        8. Power Supply
           1. 100 – 240 VAC, 4 A, 50/60 Hz
           2. Max. power consumption: 150 VA
  3. Certifications
     1. CE compliant
     2. UL certified
  4. Environmental Requirements
     1. Operational Criteria
        1. Operating temperature: 10 to 30 °C ±4 °C deviation (50 to 86 °F ±7.2 °F deviation)
        2. Relative humidity: 5 to 95 %, non-condensing
        3. Reagent temperature: keep between 10 to 30 °C (50 to 86 °F)
  5. Warranty
     1. Warranted from manufacturer defects for two years (Europe) or one year (all other geographies) from date of shipment.
  6. Maintenance and Service
     1. Unscheduled Maintenance
        1. Check and clean the analyzer components, depending on cleanliness of the sample
     2. Scheduled Maintenance / preventative
        1. Monthly
           1. Reagents refill; validation and / or calibration
        2. Quarterly
           1. Pump tubing replacement
        3. Annually
           1. Calibration
           2. Replacement of all tubing
           3. Replacement of valves and pistons

1. PRODUCTS
   1. Manufacturer
      1. Hach
      2. EZ5000 Series Alkalinity Analyzer
   2. Manufactured Unit
      1. The Alkalinity analyzer consists of a microprocessor controlled volumetric analyzer designed to monitor Alkalinity (optional: Hardness) semi-continuously in a sample stream. Automatic cleaning and validation are available.
   3. Equipment
      1. Online Analyzer
         1. Utilizes volumetric measurement with a high-precision micropumps for reagent dosing.
         2. With automatic validation, priming and cleaning
         3. If chosen, the instrument can monitor up to 8 sample streams.
         4. Plug-and-play airtight fittings for connecting reagent tubing
         5. High precision micropumps for reagent dosing
      2. Controller
         1. Industrial panel PC with 5.7” TFT color display, compact flash memory, modular compact I/O system
         2. User interface with different user levels (Automatic, User Level 1, User Level 2, Administrator)
         3. If chosen, includes capability to communicate measurements via 4-20 mA outputs, Modbus TCP/IP, Modbus RS485 or RS232.
      3. Reagents and Standards
         1. The analyzer shall use quick connect reagent containers with pre-installed tubing.
         2. Reagents and standards shall be prepared according instructions on the Method + Reagent sheet.
   4. Components
      1. Analytical instrument  
         To deliver:
         1. Alkalinity Analyzer as selected in section 1.1.A.
         2. Wall-mount bracket
         3. Reagent containers
         4. User Manual
         5. Method + Reagent Sheet
      2. Dimensions: refer to analyzer drawings
      3. Weight: 25 kg (55 lb)
   5. Instrument Options,

\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*

Must be added to instrument at time of order.

\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*

Number of sample streams (fill in, select up to 8):

[ ] sample streams

Outputs (select / fill in one)

[ ]x 4-20 mA Outputs (fill in, select up to 8)

[ ]x 4-20 mA Outputs and Modbus RS485 (fill in, select up to 4)

[ ]x 4-20 mA Outputs and Modbus TC/IP (fill in, select up to 4)

[ ] Modbus RS485

[ ] Modbus TCP/IP

[ ] RS232

* 1. Instrument Accessories

\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*

Select as many as required

\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*\*

[ ] External dilution unit

[ ] Universal Filtration System, pore size 50 µm

[ ] Universal Filtration System, pore size 100 µm

[ ] Heavy-Duty Filtration System, pore size 50 µm

[ ] Heavy-Duty Filtration System, pore size 100 µm

[ ] Table Stand

[ ] Floor Stand

1. EXECUTION
   1. Preparation
      * 1. Mounting
           1. As shown on the drawings
        2. Inlet and outlet connection sizes
           1. As shown on the drawings
        3. Sample Flow Rate
           1. 100 to 300 mL/minute
        4. Sample Pressure
           1. By external overflow vessel
        5. Sample Temperature
           1. 10 to 30 °C ±4 °C deviation (50 to 86 °F ±7.2 °F deviation)
        6. Other sample requirements
           1. Maximum particle size 100 µm, <0.1 g/L suspended solids, turbidity <50 NTU
        7. Instrument air
           1. Dry and oil free according to ISA-S7.0.01-1996 quality standard for instrument air
        8. Demineralized water
           1. For rinsing and/or dilution
        9. Drain
           1. Atmospheric pressure, vented, min. ø 64 mm
        10. Earth connection
            1. Dry and clean earth pole with low impedance (> 1 Ohm) using an earth cable of > 2.5 mm2
        11. Power supply
   2. Installation
      1. Install analyzer following transmittal drawings and instrument user manual.
   3. Manufacturer’s Service and Start-Up
      1. Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.
      2. Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.
      3. Items A and B are to be performed by manufacturer’s factory-trained service personnel. Field service and factory repair by personnel not employed by the manufacturer is not allowed.
      4. Use of manufacturer’s service parts is required. Third-party parts are not approved for use.

END OF SECTION

DOC353.53.30549.Jul20