Single or multi-channel Online Total Organic Carbon Analyser

Analyser for measuring Total Organic Carbon, Total Carbon, Total Inorganic Carbon, Biological Oxygen Demand and/ or Chemical Oxygen Demand via correlation, and Volatile Organic Carbon via calculation. The analyser is designed for measurement of these parameters in water. The Total Organic Carbon measurement method is infrared measurement of Carbon Dioxide after oxidation. The oxidation method is Two Stage Advanced Oxidation which utilises the Hydroxyl radical and a manganese catalyst.

The analyser is capable of measuring Total Organic Carbon every 6.5 minutes. The analyser can be configured to measure any combination parameters in up to 6 process streams. The sample sequencer is internal to the analyser. The analyser is self-cleaning. The analyser uses the internal sample pump to bring the sample into the reactor. The analyser reactor is self-cleaning. The analyser has internal sensors and utilizes self-diagnostics to monitor analyzer status and produces alerts when maintenance is required.

Technical Data

Measurement method

Infrared measurement of CO2 after oxidation (DIN EN 1484:1997-08, ISO 8245:1999-03, EPA 415.1)

Innovative Two-Stage Advanced Oxidation Process (TSAO) using Hydroxyl Radicals 0 - 100 mg/L C, 0 - 250 mg/L C, 0 - 1000 mg/L C, 0 - 2000 mg/L C, 0 - 10000 mg/L C,

0 - 20000 mg/L C

Automatic or Manual Range Selection 1, 2, 4 and 6 streams

±3% of reading or ±0.3 mg/L, whichever is greater in 0 - 100 mg/L C range

±3% of reading or ±0.45 mg/L, whichever is greater in 0 - 250 mg/L C range

±3% of reading or ±1.5 mg/L, whichever is greater in 0 - 1000 mg/L C range

±3% of reading or ±6 mg/L, whichever is greater in 0 - 2000 mg/L C range

±3% of reading or ±30 mg/L, whichever is greater in 0 - 10000 mg/L C range

±3% of reading or ±30 mg/L, whichever is greater in 0 - 20000 mg/L C range From 6.5 minutes, depending on range and application

1 programmable 4-20 mA analog output signal channel fitted as standard

3 freely programmable system relays Up to 30 %

Oxidation method

Measuring range

Range selection

Multi-Stream Repeatability

Cycle time Output

Digital outputs

Permissible Chloride

range

Communication: digital

Modbus RTU, Modbus TCP/IP & Profibus (when the Profibus option is selected, the digital output signals are sent through the Profibus converter with its specific communication protocol)

Except for Zone 1 certification then Modbus RTU, Modbus TCP/IP & Modbus TCP/IP

Redundant is available

IP44; optional IP54 with air purge

Enclosure waterproof rating

EExp / Hazardous

Location

Sample inlet temperature Ambient temperature

Certification options are available to European Standards, (ATEX Zone 1, Zone 2), North

American Standards (Class I Division 2) and IECEx Zone 1 2 - 60 °C

5 - 40 °C

Air conditioning and heating options are available. 5 - 85 % (non-condensing)

Typically ambient

Up to 2 mm, soft particulates

Humidity

Sample inlet pressure Particle size

PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Data storage

Previous 9999 analysis data on screen in the microcontroller memory and storage of data archive for the lifetime of the analyser in the SD/MMC card.

Previous 99 fault data on screen in the microcontroller memory and storage of fault data archive for the lifetime of the analyser in the SD/MMC card.

High contrast 40 character x 16 line backlit LCD with CFL backlight 115 V AC / 230 V AC

Display

Power requirements (Voltage)

Power requirements (Hz) Service interval Dimensions (H x W x D) Weight

50/60 Hz

6 months service intervals 1250 mm x 750 mm x 320 mm

90 - 120 kg

Enclosure weight may change depending on system optional features. TOC, TIC, TC, VOC, after correlation COD, BOD, TOG

Parameter

Scope of delivery

[ ] (delete what is not needed)

Analytical Instrument

[ ] Online TOC Analyser, 0 - 100 mg/L C, 0 - 1000 mg/L C, 0 - 10000 mg/L C

[ ] Online TOC Analyser, IP54 Purge Ready, 0 - 100 mg/L C, 0 - 1000 mg/L C, 0 - 10000 mg/L C

[ ] Online TOC Analyser, ATEX Compliant, EExp Zone 2, 0 - 100 mg/L C, 0 - 1000 mg/L C, 0 - 10000 mg/L C

[ ] Online TOC Analyser, ETL Certified, EExp Class 1 Div 2, 0 - 100 mg/L C, 0 - 1000 mg/L C, 0 - 10000 mg/L C

[ ] Online TOC Analyser, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

[ ] Online TOC Analyzer, IP54 Purge Ready, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

[ ] Online TOC Analyser, ATEX Compliant, EExp Zone 2, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

[ ] Online TOC Analyser, ETL Certified, EExp Class 1 Div 2, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

Process Streams/Channels

[ ] 1 channel [ ] 2 channels [ ] 3 channels [ ] 4 channels [ ] 5 channels [ ] 6 channels

Accessories and options

[ ] VOC mode used for measurement of Volatile Organic Carbon in the sample [ ] Remote activation of Manual/Calibration port

[ ] Profibus

[ ] Modbus RTU

[ ] Modbus TCP/IP

[ ] Sample Sensor to detect presence of the sample and evaluate the quantity of air bubbles in the sample

[ ] Sample overflow chamber

[ ] Vacuum Sampler [ ] Venturi Sampler

[ ] Flow-through sampling chamber with sand trap [ ] Sampling chamber with sand trap

[ ] Compressor

[ ] Filter pack for removal of water, dirt, and oil from compressed air (if using instrument air supply)

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:

Hach

Product:

BioTector B7000i Online TOC Analyser