Digital Online Process ORP Sensor

Online process sensor for potentiometric measurement of the oxidation-reduction potential (ORP or Redox Potential) of water. The ORP technology uses a process electrode and reference electrode to measure the ORP differentially with respect to a third ground electrode. The sensor connects to a universal digital controller.

A double junction salt bridge creates a barrier to contamination which minimizes the dilution of the internal standard cell solution. A replaceable salt bridge holds a large volume of buffer to protect the reference electrode from harsh process conditions. The salt bridge threads onto the end of the sensor if replacement is needed.

Encapsulated construction protects the sensor’s built-in preamp from moisture and humidity. A preamp in the ORP analogue sensor produces a strong enough signal, enabling the sensor to be located up to 1000 m from the controller.

Includes the capability to remotely monitor sensors on any browser-enabled device and present diagnostics on the overall health of the measurements (on predictive diagnostics-enabled sensors), as well as upcoming and required maintenance - reducing user risk and downtime.

Technical Data

Measuring range Accuracy Sensitivity Repeatability

Temperature accuracy

Drift Operating

temperature range

-1500 mV - 1500 mV

± 5 mV

± 0.5 mV

± 2 mV

± 0.5 °C

2 mV per 24 hours, non-cumulative

Digital Sensor: -5 - 70 °C

Analogue Sensor with Digital Gateway: 5 - 105 °C Immersion Sensor: 0 - 50 °C

max. 3 m/s

Digital: 6.9 bar at 70 °C

Analogue: 6.9 bar at 105 °C

NTC 300 Ω thermistor for temperature readout, not for temperature compensation

100 m, maximum

1000 m, maximum when used with a termination box

10 m PUR (Polyurethane) 4-conductor with one shield, rated to 105 °C

Stainless steel, Ryton, glass, titanium, FKM/FPM o-ring 4 - 70 °C

0 - 95 % relative humidity (non-condensing)

One point manual

Submersible to 107 m/1050 kPa

Modbus

0.870 kg

Flow

Sensor pressure / Temperature limit

Temperature sensor

Transmission distance

Cable length

Wetted materials Storage conditions

Calibration method Immersion depth Communication

Weight

PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Scope of delivery

Sensor complete with a 10 m cable and manual

Electrode Type

[ ] (Select at least 1 option)

[ ] General purpose platinum electrode for most applications

[ ] Special purpose gold electrode for applications containing zinc, cadmium, cyanide, or nickel

Sensor Body Material

[ ] (Select at least 1 option)

[ ] PEEK body (for chemical compatibility with most process solutions) [ ] Ryton body (for less aggressive solutions)

[ ] Stainless steel body (for immersion applications).

Sensor Mounting Options

[ ] (Select at least 1 option)

[ ] Convertible [ ] Insertion

[ ] Immersion [ ] Sanitary

Accessories

[ ] (delete what is not needed)

[ ] Universal Controller

[ ] 200 mV ORP solution, 500 mL [ ] 200 mV ORP solution, 1 L

[ ] 200 mV ORP solution, 3.5 L

[ ] 600 mV ORP solution, 500 mL [ ] 600 mV ORP solution, 1 L

[ ] 600 mV ORP solution, 3.5 L

[ ] Standard cell solution, concentrated pH 7.0 buffer (equi-transferrant), 500 mL [ ] PEEK salt bridge, Kynar (PVDF) outer-junction

[ ] Digital extension cable, 7.7 m [ ] Digital extension cable, 15 m [ ] Digital extension cable, 31 m

[ ] Pole mount assembly for 1” NPT sensors [ ] CPVC union mount for pHD sensor

[ ] Float mount assembly for 1” NPT sensors

[ ] CPVC flow-through mount assembly hardware for pHD differential sensors [ ] Insertion mounting, digital, 1½” CPVC

[ ] Self-contained air blast cleaning systems (115 V AC) for pHD

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:

Hach

Product:

pHD sc Digital Online Process ORP Sensor