PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Mono- or multi-channel Online Sodium Analyser

Mono-channel or multi-channel online analyser for low-level sodium measurement based on a direct potentiometric technique using a highly sensitive sodium glass electrode. The analyser uses an auto-adapted rinsing sequence with a cycle time of 10 or more minutes. To maintain optimum response time, even in systems of continuous low sodium concentration, the analyser provides automatic electrode reactivation. Reactivation uses non-hazardous chemicals and eliminates the need for manual reactivation or electrode etching.

The fully automatic analyser calibration is based on the “known addition” technique that avoids risk of contamination or human error. Grab sample feature allows the use of standards to check operation or measurement of a one-off process sample to reduce laboratory time. A manual sample (250 mL) can be introduced without disconnecting any tubes. After sampling, the instrument automatically returns to online monitoring.

Technical Data

Mono-channel Sodium Analyser

Measuring range	0 - 10.000 ppb freely programmable;

0 - 200 ppm with K-kit option

Response time	180 s (t=90%)

Calibration method	Manual: 1 or 2 points

Optional: Automatic with known addition

Number of sample	1

streams

Analogue outputs	4 × (0 or 4) /20 mA [650 Ohms] / linear and dual / Smart

Options	Conditioning kit (K-kit), automatic calibration, filtration system, wall enclosure

Multi-channel Sodium Analyser

Measuring range	0 - 10.000 ppb freely programmable;

0 - 200 ppm with K-kit option

Response time	1 cycle, minimum 10 minutes

Calibration method	Automatic with known addition: 3 points

Manual: 1 or 2 points

Number of sample streams

1, 2 or 4

Analogue outputs	6 x 0/4 to 20 mA  [800 Ohms] linear, dual or logarithm with event indication

Options	Conditioning kit (K-kit), static heat exchanger, filtration system, wall enclosure

Sodium Analyser (mono-/multi-channel)

Sample conditioner	Di-isopropylamine (1 L/100 days), recommended or

NH3 (2.5 L/90 days)

Repeatability	DIPA: < 0.02 ppb or 1.5 % of reading (whichever is greater), within 10 °C variation

NH3: < 0.1 ppb or 1.5% reading (whichever is greater), within 10 °C variation

Detection limit	DIPA: 0.01 ppb

NH3: 2 ppb

Max. concentration of suspended solids in sample

<10 ppm / <2 NTU, no oil, no grease. For boiler sample type, install approx. 100 µm filter

Acidity	Standard: <50 ppm

Using Kit K: <250 ppm (equivalent CaCO3)

pH Range	6 - 10 pH

Cationic application (using K-kit): 2 - 10 pH

Sample temperature	5 - 45 °C Ambient temperature	5 - 50 °C Sample pressure	0.2 - 6 bar

Flow rate	5 L/h during sampling phase

Inlet	Simple fittings for 6 mm OD tubing or ¼” OD in PE-low density. ¼” OD in PHED-PTFE-SS

as option

Outlet	Barbed stem for 12 mm (½”) ID hose

Power requirements

(Voltage)

Power requirements

(Hz)

100 - 240 V AC

50 - 60 Hz

Protection rating	Transmitter: IP 65 (NEMA 4)

Panel: IP50 (dust protection)

Enclosure: IP54 (splash water proof)

European standards for EMC

EN 61326-1:2006, EN 61010-1:2001

International standards   U. L. and GOST Metrology

Display	Digital, 75 x 75 mm graphic with LED backlighting

Concentration, trend curves, diagnostics, alarm status, calibration constants, historical data

User interface	Menu driven operation and clear messages in 5 languages

Optional outputs	Modbus RS485, Profibus DP

Relay output	30 V DC, 0.5 A maximum

Number of relays	4 programmable (concentration, temperature, minimum sample flow rate)

Material	Panel: ABS with stainless steel frame

Enclosure: ABS

Dimensions	Panel: 850 mm x 450 mm x 252.5 mm (H x W x D)

Enclosure: 850 mm x 450 mm x 331.5 mm (H x W x D)

Weight	Panel: 18 kg; Enclosure: 23 kg; Full canisters: 2 kg

Maintenance interval	Every 100 days: refill electrolyte, reagents, and calibration solution (using DIPA)

Scope of delivery

[ ] (delete what is not needed) Analytical Instrument

[ ] Mono-channel Online Sodium Analyser

[ ] Multi-channel Online Sodium Analyser, 1-channel version [ ] Multi-channel Online Sodium Analyser, 2-channel version [ ] Multi-channel Online Sodium Analyser, 3-channel version [ ] Multi-channel Online Sodium Analyser, 4-channel version

Communication

[ ] Standard

[ ] With RS485 (Panel mount only)

[ ] With Profibus (Wall mount only)

Mounting style

[ ] Panel mount

[ ] Wall mount

Conditioning kit

[ ] Standard (0.01-10,000 ppb) [ ] Kit K (1 ppb - 200 ppm)

Accessories

[ ] Sample filter, 100 µm, metric fittings

[ ] Sample filter, 100 µm, imperial fittings

[ ] Static heat exchanger system, imperial fittings

Mounting Hardware

[ ] One year spare part kit

[ ] Kit of canisters

[ ] Bottle cap adapter

[ ] Monoethylamine bottle cap adapter

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Hach

Product:	Polymetron 9240/9245 Online Sodium Analyser