UV probe for dissolved organic substances in water

Stainless steel process probe for continuous, reagent-free measurement of dissolved organic substances in water, wastewater, surface water, process water and solids-free landfill leachate. The method of measuring UV absorbance and percent transmittance is determining the Spectral Absorption Coefficient (SAC) at a wavelength of 254 nm, using the

2-beam technology. The SAC254 is a measure of the organic content of the fluid.

The integrated cleaning system uses wiper technology. It allows the probe to measure dissolved organic substances in

media with Suspended Solid contents. The probe requires a universal digital controller.

Includes the capability to remotely monitor sensors on any browser-enabled device and present diagnostics on the overall health of the measurements (on predictive diagnostics-enabled sensors), as well as upcoming and required maintenance - reducing user risk and downtime.

Technical Data

UV probe for organic load, immersion (version 1)

Path Length

Measuring range

1 mm

SAC254: 2 - 3000 m-1 (can be calibrated to other parameters (e.g. TOC, COD, BOD) depending on the application)

Aeration basin

Application

UV probe for organic load, immersion or bypass (version 2)

Path Length

Measuring range

2 mm

SAC254: 0 - 1500 m-1 (can be calibrated to other parameters (e.g. TOC, COD, BOD) depending on the application)

Aeration basin

Application

UV probe for organic load, immersion or bypass (version 3)

Path Length

Measuring range

5 mm

SAC254: 0.1 - 600 m-1 (can be calibrated to other parameters (e.g. TOC, COD, BOD) depending on the application)

Outlet; Drinking water; Surface water

Application

UV probe for organic load, immersion or bypass (version 4)

Path Length

Measuring range

50 mm

SAC254: 0.01 - 60 m-1 (can be calibrated to other parameters (e.g. TOC, COD, BOD) depending on the application)

Drinking water; Clear water

Application

UV probe for organic load (in general, all versions)

Measurement method

UV absorption measurement (unique 2-beam technique), reagent-free

SAC254 in accordance with DIN 38404 C3 550 nm

≥ 1 min

Compensation

Measuring interval

PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Cable length

10 m fixed cable at sensor

Optional extension cables available in 5, 10, 15, 20, 30, 50 m Total maximum length: 60 m

0.5 bar

2 - 40 °C

Pressure range

Operating temperature range

Diameter Length Weight

Service interval

Maintenance requirements

Controller

compatibility

70 mm

333 mm

3.6 kg

6 months

1 h/month in standard applications

Universal Controller, Multi-Parameter Universal Controller

Scope of delivery

[delete what is not needed]

Analytical Instrument

[ ] UV probe with 1 mm path length, SAC254: 2 - 3000 m-1 (version 1, immersion only)

[ ] UV probe with 2 mm path length, SAC254: 0 - 1500 m-1 (version 2, immersion or bypass) [ ] UV probe with 5 mm path length, SAC254: 0.1 - 600 m-1 (version 3, immersion or bypass)

[ ] UV probe with 50 mm path length, SAC254: 0.01 - 60 m-1 (version 4, immersion or bypass)

Mounting Hardware

[ ] Mounting assembly kit “Rim Mounting” (immersion application)

[ ] Mounting assembly kit “Flow-Through” for 2 mm probe (bypass application) [ ] Mounting assembly kit “Flow-Through” for 5 mm probe (bypass application)

[ ] Mounting assembly kit “Flow-Through” for 50 mm probe (bypass application)

Accessories

[ ] Digital extension cable	(select 5, 10, 15, 20, 30 or 50 m)

[ ] Set of wiper blades	(select 1, 2, 5, 50 mm)

Service

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:

Hach

Product:

UVAS plus sc UV probe for dissolved organic substances in water