Process turbidity probe with predictive diagnostics

Process probe for continuous monitoring and measurement of turbidity or turbidity and suspended solids. The method of measuring turbidity and suspended solids will be with a dual-beam infrared/scattered light photometer (turbidity measurement in accordance with DIN ISO EN 27027, Total Suspended Solids (TSS) measurement equivalent to DIN 38414). Measurement is independent of coloration. The sensor is factory calibrated and needs no calibration prior to use. It allows for individual calibration up to 5 calibration points. The sensor requires a universal digital controller.

The sensor has built in predictive diagnostics capability which increases available operation time. The predictive diagnostics produces alerts for upcoming maintenance tasks by monitoring the instrument’s internal components and tracking service requirements. When connected to a properly equipped controller the controller will display status indicators for each sensor with predictive diagnostics capability.

Includes the capability to remotely monitor sensors on any browser-enabled device and present diagnostics on the overall health of the measurements (on predictive diagnostics-enabled sensors), as well as upcoming and required maintenance - reducing user risk and downtime.

Technical Data

Process Turbidity Probe for immersion in open tanks, Version 1

Parameter

Measuring range turbidity

Measuring range TSS

Turbidity

0.001 - 4000 NTU

-

Process Turbidity Probe for immersion in open tanks, Version 2

Parameter

Measuring range turbidity

Measuring range TSS

Suspended Solids, Turbidity

0.001 - 4000 NTU

0.001 - 50 g/L / 0.001 - 50,000 mg/L

Process Turbidity Probe for immersion in open tanks, Version 3

Parameter

Measuring range turbidity

Measuring range TSS

High Range Suspended Solids, Turbidity

0.001 - 4000 NTU

0,001 - 500 g/L / 0.001 - 500,000 mg/L

(the upper range is dependent on the sample property)

Process Turbidity Probe for insertion in pipes, Version 4

Parameter

Measuring range turbidity

Measuring range TSS

Suspended Solids, Turbidity

0.001 - 4000 NTU

0.001 - 50 g/L / 0.001 - 50,000 mg/L

Process Turbidity Probe for insertion in pipes, Version 5

Parameter

Measuring range turbidity

Measuring range TSS

High Range Suspended Solids, Turbidity

0.001 - 4000 NTU

0.001 - 500 g/L / 0.001 - 500,000 mg/L

(the upper range is dependent on the sample property)

PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Process Turbidity Probe (in general for all Versions)

Units

Turbidity: User selectable - NTU, FNU, or TE/F

Suspended Solids: User selectable - g/L, mg/L, ppm, or % solids

Accuracy

Repeatability

Turbidity: < 1% TSS content: < 3% 1 - 300 s adjustable

Turbidity: Formazin or Stablcal Standard (at 800 NTU).

Suspended Solids: Sample specific, based on gravimetric TSS analysis with a correction

factor procedure.

CE

Max. 3 m/s (the presence of air bubbles affects the measurement)

0 - 40 °C

Response time

Calibration method

Certifications Flow

Operating temperature range

Pressure range

PVC: 1 bar or 10 m (only for version 1, version 2 and version 3)

Stainless steel insertion sensor: 6 bar or 60 m (not available for version 1)

Optics carrier and sleeve: stainless steel 1.4571; Wiper arm: stainless steel 1.4581; Wiper rubber: silicone rubber (standard) Optional: FKM/FPM (LZX578); Wiper shaft: stainless steel 1.4571

Immersion:

Stainless steel: 1.38 kg PVC: 0.52 kg

Insertion:

Stainless steel: 2.4 kg

10 m (optional extension cables available)

Material

Weight

Cable length

Scope of delivery

[ ] (delete what is not needed)

Analytical Instrument

[ ] Process turbidity probe for immersion in open tanks, Version 1 [ ] Process turbidity probe for immersion in open tanks, Version 2 [ ] Process turbidity probe for immersion in open tanks, Version 3 [ ] Process turbidity probe for insertion in pipes, Version 4

[ ] Process turbidity probe for insertion in pipes, Version 5

Cleaning option

[ ] With automatic wiper cleaning (wiper made of silicon) [ ] Without cleaning option

Housing material

[ ] Stainless steel housing (not available for Version 1) [ ] PVC housing (for immersion only)

Accessories

[ ] Digital extension cable,	(select 5/10/15/20/30/50 m)

[ ] Set of wiper blades for probe, made of silicone for standard applications (spare parts) [ ] Set of wiper blades for probe, made of FKM/FPM for media containing oil

[ ] Set of seals only for probes with wiper [ ] Crank-handle (for immersion only)

[ ] O-Ring for inline fitting (for immersion only)

Mounting Hardware

[ ] Mounting hardware as follows:

[ ] Rim mounting, stainless steel, with 90° adapter, 10 cm base, 2 m pole (for immersion only) [ ] Rim mounting, stainless steel, with 90° adapter, 24 cm base, 2 m pole (for immersion only) [ ] Welding neck flange made of stainless steel (for insertion only)

[ ] Welding neck flange made of C steel (for insertion only)

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:

Hach

Product:

Solitax sc Process turbidity probe with Prognosys predictive diagnostics