EMBED Word.Picture.8

PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Automatic Refrigerated Sampler

An automatic sampler to collect and store unbiased representative water samples for laboratory analysis. A refrigerated sampler is optimised for indoor applications where temperature control is required for sample preservation.  Equipped with a peristaltic pump with spring-mounted rollers and a standard liquid detect sensor for sample volume accuracy. A USB connection and an auxiliary port for pulse or 4-20 inputs are standard. Factory-installed options include sensor ports for integrated water quality, level or flow, a port for Modbus RS485 communication or a rain sensor and an alternate non- contact liquid detect sensor. Using appropriate modules, inputs and outputs can be extended for a single high-voltage relay or a comprehensive set of analogue outputs or analogue and digital outputs and relays. Refrigerated sampler base is available in black vinyl or stainless steel finish.

# Technical Data

| Sampling Features |  |
| --- | --- |
| Programs | Dual Programs: Up to 2 sample programs can be run sequentially, in parallel, or according to day of week scheduling; enabling a single sampler to function like multiple samplers |
| Sampling mode | Pacing: Time weighted, Flow weighted, Time table, Flow table, Event Distribution: Single bottle composite, multi-bottle composite, multi-bottle discrete, bottles per sample, samples per bottle or a combination of bottles per sample and samples per bottle |
| Operating mode | Continuous or non-continuous |
| Status messages | Communicates what program is running, if there are any missed samples, when the next sample will be taken, how many samples remain, number of logged channels, time of last measurement, memory available, number of active channels, if alarms were triggered, when alarms were triggered, active sensors and cabinet temperature |
| Alarm | Configurable alarms that show on status screen and are recorded in diagnostics alarm logs. Alarms can be set for system diagnostics and logging such as program end, sample complete, missed samples and full bottle. Channel alarms are setpoint alarms for the recorded measurements (channels), such as pH, level and power supply voltage. |
| Manual sampling | Initiates a sample collection independent of program in progress |
| Automatic shutdown | Multiple Bottle Mode: After complete revolution of distributor arm (unless Continuous Mode is selected) Composite Mode: After preset number of samples have been delivered to composite container, from 1 to 999 samples, or upon full container. |
| Sample volume | Programmable in 10 mL increments from 10 to 10,000 mL |
| Sample interval | Selectable in single increments from 1 to 9999 flow pulses or 1 to 9999 minutes in one minute increments |
| Sample trigger | When equipped with flow sensor or pH/temperature sensor or peripheral monitoring options, sampling can be triggered upon an upset condition when field selectable limits are exceeded |
| Data logging | SAMPLE HISTORY - Stores up to 4000 entries for sample time stamp, bottle number and sample status (success, bottle full, rinse error, user abort, distributor error, pump fault, purge fail, sample timeout, power fail and low main battery). MEASUREMENTS - Stores up to 325,000 entries for selected measurement channels in accordance with the selected logging interval. EVENTS - Ability to store up to 2000 entries in Sample History logging. Records Power On, Power Fail, Firmware Updated, Pump Fault, Distributor Arm Error, Low Memory Battery, Low Main Battery, User On, User Off, Program Started, Program Resumed, Program Halted, Program Completed, Grab Sample, Tube Change Required, Sensor Communication Errors, Cooling Failed, Heating Failed, Thermal Error Corrected. |
| Diagnostics | View event and alarm logs as well as maintenance diagnostics |

| Refrigerated Sampler |  |
| --- | --- |
| Housing material | 22 Gauge Steel (optional Stainless Steel) with Vinyl laminate over-coating (Refrigerator) |
| Refrigeration components and plumbing | Corrosion protected with conformal all exposed copper tubing is insulated to avoid sweating and condensation |
| Cooling system | 1/7 HP, 75 Watt, 400 BTU/hr compressor, 120 CFM condenser fan, three-sided wraparound plate type evaporator, rigid foam insulation, air sensing thermostat capable of maintaining sample liquid at 4 °C in ambient temperature to 49 °C maximum; accurate to ±0.8 °C, magnetic door seal |
| Sample container | SINGLE BOTTLE: - 10 L Glass or Polyethylene, or 21 L Polyethylene MULITIPLE BOTTLES: - 10 L Polyethylene or Glass, (4) 10 L Polyethylene or Glass, (8) 2.3 L Polyethylene or 1.9 L Glass, (12) 2 L Polyethylene, (24) 1 L Polyethylene or 350 mL Glass |
| Temperature | Operating: 0 - 50 °C Storage: -40 - 60 °C |
| Power requirements (Voltage) | 115/230 V AC |
| Power requirements (Hz) | 50/60 Hz |
| Dimensions | 61 cm x 61 cm x 112 cm |
| Weight | 63 kg |
| Controller |  |
| Housing material | PC/ABS blend, NEMA 4X, 6, IP68, corrosion and ice resistant (Controller) |
| Display | 1/4 VGA, colour; self-prompting/menu-driven program |
| User interface | Membrane switch keypad with 2 multiple function soft keys |
| Languages user interface | en, fr, es, it, de, pt, zh, tr, hu, cz, pl, ro, hr, el, sl, sk, fi, ru, ja, ko |
| Lock function | Access code protection prevents tampering |
| Memory | Sample history: 4000 records; Data log: 325,000 records; Event log: 2000 records |
| Communication capabilities | USB and optional RS485 (Modbus) |
| Inputs | One 0/4-20 mA input for flow pacing |
| Certifications | CE, UL |
| Sample Pump and Strainer |  |
| Sample pump | High-speed peristaltic, dual roller, with 0.95 ID x 1.6 OD cm (3/8” ID x 5/8” OD) pump tube |
| Vertical lift | 8.5 m using 8.8 m maximum of 3/8” Vinyl intake tube at sea level at 20 to 25 °C |
| Tubing | Pump tubing: 9.5 mm ID x 15.9 mm OD Silicone Intake tubing: 1.0 - 4.75 m minimum length, 1/4” or 3/8” ID Vinyl or 3/8” ID PTFE-lined Polyethylene with protective outer cover (black or clear) |
| Sample volume repeatability | Typical: ±5% of 200 mL sample volume with: 4.6 m vertical lift, 4.9 m of 3/8” Vinyl intake tube, single bottle, full bottle shut-off at room temperature and 1524 m elevation |
| Sample volume accuracy | Typical: ±5% of 200 mL sample volume with: 4.6 m vertical lift, 4.9 m of 3/8” Vinyl intake tube, single bottle, full bottle shut-off at room temperature and 1524 m elevation |
| Sample transport velocity | 0.9 m/s with 4.6 m vertical lift, 4.9 m of 3/8” Vinyl intake tubing, 21 °C and 1524 m elevation |
| Pump flow rate | 4.8 L/min at 1 m vertical lift with 3/8” intake tube typical |
| Internal clock | ±1 second per day at 25 °C |

| Intake | Strainers: Choice of PTFE and 316 stainless steel construction, or all 316 stainless steel in standard size, high velocity, and low profile for shallow depth applications Purge: Air purged automatically before and after each sample; duration automatically compensates for varying intake line lengths Rinse: Intake line automatically rinsed with source liquid prior to each sample, from 1 to 3 rinses Retries or Fault: Sample collection cycle automatically repeated from 1 to 3 times if sample not obtained on initial attempt |
| --- | --- |

Scope of delivery

[ ] (delete what is not needed)

[ ] Basic Controller (includes 4-20 aux port and USB port) [ ] Controller with sensor ports

[ ] Controller with Rain/RS485 port

[ ] Controller with sensor ports and Rain/RS485 port

Refrigerated base

[ ] Vinyl base, 115 V [ ] Vinyl base, 230 V

[ ] Stainless steel base, 115 V [ ] Stainless steel base, 230 V

Power Back up Option

[ ] AC power backup,16”, 3 pin

Bottle Sets

[ ] Composite 21 L Polyethylene bottle kit [ ] Composite 10 L Polyethylene bottle kit [ ] Composite 10 L Glass bottle kit

[ ] Discrete (2) 10 L Polyethylene bottle kit [ ] Discrete (4) 10 L Polyethylene bottle kit [ ] Discrete (24) 1 L Polyethylene bottle kit [ ] Discrete (24) 350 mL Glass bottle kit

Intake Tubing

[ ] Vinyl - 7.5 m, 3/8” ID x 5/8” OD tubing and PTFE/stainless steel strainer [ ] PTFE-lined Polyethylene 7.5 m, 3/8” ID x 5/8” OD with connector kit

Sensors

[ ] Non oil-filled 0-3 m range, 9 m cable with connector [ ] Oil-filled 0-3 m range, 9 m cable with connector

[ ] Digital pH sensor

[ ] Digital pH sensor for Europe only [ ] Downlooking ultrasonic sensor

[ ] Suspended downlooking ultrasonic sensor [ ] In-pipe ultrasonic sensor

Accessories

[ ] Cable assy, aux half, 7P, 1.2 m [ ] Cable assy, aux half, 7P, 7.6 m

[ ] Universal junction box for 4-20 mA [ ] Input/Output 1 relay

[ ] Input/Output digital outputs/relays, analogue inputs

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Hach

Product:	AS950 Automatic Refrigerated Sampler