CSI Specification – Water Quality Monitoring Panel (WQMP) and Controller

- General Description

  - The Water Quality Monitoring Panel shall be a single panel with instruments pre-mounted, wired and plumbed.

    - The back plane of the panel shall be not greater than 28.6” wide by 40” long and made of corrosion resistant material.

    - The panel shall be fully assembled with all instruments, controller, mounted, plumbed, and wired prior to shipment.

    - The sample and drain system for the panel shall accommodate flows of 150 to 800 ml/min.

  - Compliance met with this assembly.

    - Electromagnetic Compatibility (EMC)

- US Title 47 CFR - Part 15, Subpart B

- Canada ISED ICES-003, Issue 7

    - Safety (Controller & Panel Fluidics):

- UL 61010-1

- CSA C22.2 No. 61010-1-12

    - Environmental Compliance for the U.S. and Canada markets only

- Prohibition of Certain Toxic Substances Regulations, 2012 (the 2016

Amendments) Canada Prohibition of Certain Toxic Substances

- Section 65 of the Canadian Environmental Protection Act, 1999

(CEPA) Canada Virtual Elimination List

- (EU) 2019/1021 EU POP: Persistent Organic Pollutants Regulation

- Long-Chain Perfluoroalkyl Carboxylate and Perfluoroalkyl Sulfonate

Chemical Substances; Significant New Use Rule (July 27, 2020) US

TSCA PFAS Surface Coating SNUR

- Regulation of Persistent Bio accumulative and Toxic Chemicals

under TSCA section 6(h) (December 2020) US TSCA Section 6

Rules

- Analytical Instruments.

All analytical instruments installed on the panel shall be from a single manufacturer to minimize maintenance and support issues. Please find the CSI specs for attached instruments in the links below:

  - Controller:

https://cdn.bfldr.com/7FYZVWYB/at/f8j5mt6q5gzcktvckt7rtr/SC1000_CSI_Specifications.docx

https://cdn.bfldr.com/7FYZVWYB/at/gm29sbb65zh7k5tzsjwxsm/SC1000_Display_CSI_Specifications.docx

- Installation

- Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

- Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

- Manufacturer’s Service and Start UP

- Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

- Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

- Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

- Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

- Manufacturer’s Top Level Part Numbers and associated instruments to be covered

with this specification:

| Part Number | Description |
| --- | --- |
| 8828000 | WQMP, AUTO CLEAN, SC1000 |