PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Oxygen and Hydrogen measurement in Nuclear reactor coolant system (RCS) or in Waste/Off gas applications

Optical sensing of oxygen is based on the measurement of the red fluorescence of a dye/indicator illuminated with a blue light. The oxygen concentration is calculated by measuring the decay time of the fluorescence intensity. The excitation signal is modulated, the decay time is transformed into a phase-shift of the modulated fluorescence signal, which is independent of fluorescent intensity and thus of potential aging. A calibration curve links the phase shift to the oxygen partial pressure through a Stern-Volmer equation. The dissolved oxygen concentration is then calculated with Henry’s

law using the water solubility curve as a function of the temperature.

The optical technology is not sensitive to flow variation and presence of magnetite does not impact measurement accuracy. It is not interfered with by H2. Only a yearly replacement of the luminescent spot is required.

The thermal conductivity sensor to measure hydrogen uses the ability of a given gas to transfer heat. Each gas has a

different thermal conductivity. The sensor’s measuring technique is a combination of a gas diffusion membrane and a gas thermal conductivity detector. The small volume enclosed between the diffusion membrane and the thermal conductivity detector is periodically flushed with a purge gas. After each purge, the gas to be measured diffuses from the sample through the membrane, changing the thermal conductivity of the gas surrounding the detector. It is the rate of change of the thermal conductivity that allows the concentration of the gas to be calculated. The sensor only requires an annual maintenance.

The complete system consists of an optical sensor oxygen sensor, a thermal conductivity hydrogen sensor, a controller, sensor cables and flow-through chambers.

The system can be set up to provide diagnostics information and service/calibration reminder through any communication channel including 0/4-20 mA.

Technical Data

Hydrogen Thermal Conductivity Sensor, Membrane A

Applications	Waste gas (PWR), off gas (BWR), and reactor water (BWR using HWC)

Thickness (membrane)   25 µm Material	PFA Radiation Dose Limits	105 rad

Range at 25 °C	0 to 2 ppm, or 0 to 25 cc/kg, or 0 to 1.5 bar

Flow rate	220 mL/min

Accuracy, the greater

of

(Sample temperature 20 to 50 °C, within ±5 °C of calibration temperature)

Accuracy, the greater

of

(Sample temperature 0 to 50 °C, independent of calibration temperature)

±1% of reading, or

±2 ppb, or

±0.03 cc/kg, or

±1.5 bar

±3% of reading, or

±15 ppb, or

±0.18 cc/kg, or

±6 mbar

Hydrogen Thermal Conductivity Sensor, Membrane B Applications	Reactor coolant (PWR) Thickness (membrane)   25 µm

Material	ETFE

Radiation Dose Limits	108 rad

Range at 25 °C	0 to 10 ppm, or 0 to 120 cc/kg, or 0 to 6 bar

Flow rate	200 mL/min

Accuracy, the greater

of

(Sample temperature 20 to 50 °C, within ±5 °C of calibration temperature)

Accuracy, the greater

of

(Sample temperature 0 to 50 °C, independent of calibration temperature)

±1% of reading, or

±8 ppb, or

±0.1 cc/kg, or

±6 mbar

±3% of reading, or

±60 ppb, or

±0.6 cc/kg, or

±20 mbar

Hydrogen Thermal Conductivity Sensor, in general (for all membranes)

Temperature range	0 to 50 °C

Pressure range at 25

°C

0 to 20 bar / up to 170 bar for high pressure sensor models

Cycle time	17 s

Recommended purge gas

Recommended

calibration gas

Pure N2 or air

Pure H2

Signal drift (per year)	<1% of reading

Certifications	2004/108/EC - EN 61326-1

Weight	0.95 kg

Maximum distance to

Analyser

50 m

Dissolved Oxygen Sensor, Low Level

Measuring range	0 - 2000 dissolved O2 (DO)

Temperature range	Accurate from -5 to 50 °C

Resistant to temperature from -5 to 100 °C Repeatability	± 0.4 ppb or 1 % whichever is greater Reproducibility	± 0.8 ppb or 2 % whichever is greater Accuracy	± 0.8 ppb or 2 % whichever is greater Detection limit	0.6 ppb

Response time	< 30 (90%) <10 s (gas phase); <30 s (liquid phase)

Display resolution	0.1 ppb

Calibration	Single point zero calibration with standard 99.999% nitrogen (quality 5.0) or equivalent

oxygen free gas

Sample pressure	1 - 20 bar absolute

Controller

Enclosure Construction   Wall (pipe) mount: stainless steel

Panel mount: aluminum

Enclosure waterproof rating

Compliance

certifications

Wall (pipe) mount: IP65, NEMA 4x

Panel mount: IP65

EMC: EN61326-1:2006

CE: EN61010-1:2010

ETL, conforming to UL 61010-1 and CSA 22.2 No. 61010-1

Display	Colour TFT touchscreen display

Analogue outputs	3 smart 0/4-20 mA (500 ohms), programmable as linear or tri-liniear, configurable to send

diagnostics information

Relays	3 measurement alarm relays (2A to 30 VAC or 0.5 A to 50 VDC); configurable to send

diagnostics information

1 system alarm relay (2 A to 30 VAC or 0.5 A to 50 VDC)

Communication	RS485

Profibus DP

Ethernet

USB-host to download data with a USB memory stick

Data storage	Rolling buffer or store once mode for up to 1000 measurements and 1000 operator actions

Holds calibration records for last 50 calibrations

User interface	Touch screen panel displays: concentration, trend graph, diagnostics, alarm status, historical data

Dimensions (H x W x D)	Wall dimensions: 230.5 mm x 250 mm x 160 mm (9.1 in x 9.8 in x 6.3 in)

Panel dimensions: 156 mm x 220 mm x 253.5 mm (6.14 in x 8.86 in x 9.84 in)

Power	Universal 100/240 VAC @ 50/60 Hz, 25 VA

10-36 VDC, 25 W

Scope of delivery

[ ] (delete what is not needed) Analytical Instrument

[ ] Dual channel measurement system with optical oxygen sensor, hydrogen thermal conductivity sensor, and 110-230

VAC controller to be wall-mounted (default 3 m cable)

[ ] Dual channel measurement system with optical oxygen sensor, hydrogen thermal conductivity sensor, and 110-230

VAC controller to be panel-mounted (default 3 m cable)

Mounting Hardware

[ ] Standard stainless steel flow chambers with 1/4” fittings for liquid applications

[ ] Standard stainless steel flow chambers with 6 mm fittings for liquid applications

[ ] Multi parameter flow chamber in stainless steel with 6 mm fittings for gas phase applications

[ ] Multi parameter flow chamber in stainless steel with 1/4” fittings for gas phase applications

Accessories

[ ] Pressure sensor for gas phase applications (0-5 bar)

[ ] Replacement luminescent spot for oxygen sensor, low range (0-2000 ppb dissolved, 0-40 mbar) [ ] Oxygen sensor calibration device (does not include calibration gas)

[ ] Hydrogen sensor maintenance kit

[ ] Water separation filter

[ ] 5 m sensor cables

[ ] 10 m sensor cables

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Hach

Product:	Orbisphere 511 Controller, with Orbisphere K1200 and 31290 Sensors