PART 1   GENERAL

1.1	Product overview

    - The NP6000sc analyzer measures phosphate ions (PO4-P) in aqueous solutions (e.g., wastewater, process water and surface water). The analyzer is used with an SC Controller for power and operation. The measurement value on the display is shown in mg/L (or ppm) of PO4-P or PO4. The conversion formula is: PO4-P x 3.07 = PO43.

    - Two models of the analyzer are available as single or dual channel with external or integrated sample filtration systems.

    - The single channel analyser includes grab sample capability for direct correlation with laboratory reference measurements. The same feature can be used to do a verification with a standard solution.

    - Includes the capability to remotely monitor sensors on any browser-enabled device and present

diagnostics on the overall health of the measurements (on Predictive Diagnostics-enabled sensors),

as well as upcoming and required maintenance - reducing user risk and downtime.

1.2	Measurement Principle

- The reagents (and optional standard) used for the chemical analysis are installed in the analyzer enclosure. The analyzer uses pumps and valves to move the sample and reagents to the measurement cell. The phosphate ions in the sample cause a yellow color reaction in the photometer cuvette when mixed with the reagents. The color change is found with the photometer. When the measurement cycle is complete, the analyzer discards the sample through the drain line. The analyzer can automatically start cleaning intervals for all measurement ranges.

- Measurement range 1 (0.015 to 5 mg/L): The analyzer automatically calibrates the reagent offset at a set frequency (recommended: weekly).

- Measurement range 2 (0.05 to 15 mg/L) and Measurement range 3 (1 to 75 mg/L): The analyzer is calibrated at the factory. An on-site calibration is not necessary for the life of the analyzer.

- The automatic calibration is only necessary for Measurement range 1. No regular calibration is necessary for Measurement ranges 2 and 3.

B.  At the beginning of each measurement cycle the analyzer measures the sample for purposes of compensating for background color of the sample.

1.3	Alternates

- Other instruments that do not use the vanadomolybdophosphoric acid colorimetric method are not acceptable and that do not compensate for background sample color are not acceptable

- Other instruments that do not have predictive diagnostic capabilities are unacceptable

1.4	System Description

A.  Performance Requirements

      - Measurement range (depending on model):

        - 0.05 to 5 mg/L,

        - 0.05 to 15 mg/L

        - 1 to 75 mg/L

      - Lower detection limit (depending on model):

        - 0.015 mg/L,

        - 0.05 mg/L

        - 1 mg/L

      - Accuracy (depending on model):

        - 2% ±0.015 mg/L,

        - 2% ±0.05 mg/L

        - 2% ±1.0 mg/L

      - Reproducibility (depending on model):

        - 0.7% ±0.005 mg/L,

        - 2% ±0.05 mg/L

        - 2% ±1.0 mg/L

      - Response time:

Measurement range 1 (LR): Less than 10 minutes (T90), with sample preparation FX6

Measurement range 2 and 3 (MR &HR):  Less than 5 minutes (T90), with sample preparation FX6

      - Measurement interval:

Measurement range 1: 10, 15, 20 or 30 minutes, adjustable

Measurement range 2 and 3: 5, 10, 15, 20 or 30 minutes, adjustable

1.5	Certifications

- CE, UKCA, CMIM, FCC, ISED, certified to UL and CSA safety standards by TÜV

1.6	Environmental Requirements

A.  Operational Criteria

1.   Sample temperature: 4 to 40 °C (39 to 104 °F)

2.   Sample pH: 5 to 9 pH

3.   Operating temperature: -20 to 45 °C (-4 to 114 °F)

4.   Operating humidity: 95% relative humidity, non condensing

1.7	Warranty

- The products include a one-year warranty from date of shipment.

PART 2   PRODUCTS

2.1	Manufacturer

A.  Hach Company

1. Model NP6000SC Phosphate Analyzer

2.2	Manufactured Unit

A.  The NP6000SC Phosphate Analyzer consists of:

1. Enclosure Rating IP55, NEMA, UL50E 3R

Material PUR 66

2. Colorimeter

2.3	Equipment

- The analyzer cleans itself automatically.

- The Low Range analyser does an offset calibration automatically

C.  The required power supply is 230 Vac/50Hz or optional 115 Vac/50-60Hz connected to a Hach model  sc1000 multi-parameter or SC4500 controller.

D.  Data transmission is made with a data cable with the controller.

E.  Outputs include relay, current outputs, and bus interface via the controller

2.4	Components

A.  Standard equipment:

1.   Analyzer

2.   Manual and mounting brackets and small parts

3.   Reagents

B.  Dimensions: 575 × 991 × 425 mm (22.63 × 39.01 × 16.73 inch)

C.  Weight: 45 kg (99.21 Ib.) without chemicals

2.5	Accessories

A. FX6 filtration module or Hach Filtrax

B.   Hach SC1000 multi-parameter controller and the SC4500 controller

C.  Mounting kits (rail or stand)

D.   Reagents set

PART 3   EXECUTION

3.1	Preparation

A. Wall-mount indoors or outdoors.

B. Rail- and stand-mounting options available

3.2	Installation

A.  Contractor will install the sensor in strict accordance with the manufacturer’s instructions and recommendation.

B.  Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

1.   Contractor will schedule a date and time for start-up.

2.   Contractor will require the following people to be present during the start-up procedure.

  - General contractor

  - Electrical contractor

  - Hach Company factory trained engineer

  - Owner’s personnel

  - Engineer

  - Manufacturer’s Service and Start-Up

  - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

  - The contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

  - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

  - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

END OF SECTION