Single channel Online Total Organic Carbon Analyser

Analyser for measuring Total Organic Carbon, Total Inorganic Carbon, Biological Oxygen Demand and/ or Chemical Oxygen Demand via correlation. The analyser is designed for measurement of these parameters in water. The Total Organic Carbon measurement method is infrared measurement of Carbon Dioxide after oxidation. The oxidation method is Two Stage Advanced Oxidation which utilises the Hydroxyl radical and a manganese catalyst.

The analyser is capable of measuring Total Organic Carbon every 7.5 minutes. The analyser can be configured to

measure any combination parameters in the process stream. The sample sequencer is internal to the analyser. The analyser is self-cleaning. The analyser uses the internal sample pump to bring the sample into the reactor. The analyser reactor is self-cleaning. The analyser has internal sensors and utilises self-diagnostics to monitor analyser status and produces alerts when maintenance is required.

Technical Data

Parameter

Direct measurement of Total Organic Carbon, Total Inorganic Carbon

Chemical Oxygen Demand, Biological Oxygen Demand via correlation 0 - 250 mg/L C, 0 - 1000 mg/L C

Automatic or manual range selection

±3% of reading or ±0.45 mg/L, whichever is greater in 0 - 250 mg/L C range

±4% of reading or ±2 mg/L, whichever is greater in 0 - 1000 mg/L C range

Infrared measurement of CO2 after oxidation (DIN EN 1484:1997-08, ISO 8245:1999-03, EPA 415.1)

Unique Two-Stage Advanced Oxidation Process (TSAO) using Hydroxyl Radicals Single process stream and manual grab sample

Measuring range Range selection

Repeatability

Measurement method

Oxidation method

Number of sample streams

Cycle time Output

Digital outputs Particle size Automatic cleaning Data storage

Communication: digital

Typically 7 minutes, 30 seconds

4 programmable 4-20 mA analogue output signal channel fitted as standard

6 freely programmable system relays Up to 100 µm

Yes

Previous 9,999 reaction data

Modbus RTU, Modbus TCP/IP & Profibus (when the Profibus option is selected, the digital output signals are sent through the Profibus converter with its specific communication protocol)

Microcontroller with membrane keyboard

High contrast 40 character x 16 line backlit LCD with LED backlight 2 - 60 °C

5 - 45 °C

Typically ambient

5 - 85 % (non-condensing) 6 month service intervals 115/230 V AC

User interface Display

Sample inlet temperature Ambient temperature Sample inlet pressure Humidity

Service interval

Power requirements (Voltage)

Power requirements (Hz) Dimensions (H x W x D) Weight

60 Hz

750 mm x 500 mm x 320 mm

46 kg (enclosure weight may change depending on system optional features)

PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Scope of delivery

[ ] (delete what is not needed)

Analytical Instrument

[ ] Single channel Online Total Organic Carbon Analyser, 0 - 250 mg/L C

[ ] Single channel Online Total Organic Carbon Analyser, 0 - 250 mg/L C and 0 - 1000 mg/L C [ ] Single channel Online Total Organic Carbon Analyser, IP54 Purge Ready, 0 - 250 mg/L C

[ ] Single channel Online Total Organic Carbon Analyser, IP54 Purge Ready, 0 - 250 mg/L C and 0 - 1000 mg/L C

Accessories and options

[ ] Modbus [ ] Profibus [ ] Ethernet

[ ] Network Control Unit

[ ] Sample Sensor to detect presence of the sample and evaluate the quantity of air bubbles in the sample

[ ] Sample overflow chamber

[ ] Flow-through sampling chamber with sand trap [ ] Sampling chamber with sand trap

[ ] Compressor

[ ] Filter pack for removal of water, dirt, and oil from compressed air (if using instrument air supply)

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:

Hach

Product:

BioTector B3500e Online TOC Analyser