Digital Online Process pH Sensor

Online process pH sensor for measuring pH of water. The pH technology uses a process electrode and reference electrode to measure the pH differentially with respect to a third ground electrode. The sensor connects to a universal digital controller.

A double junction salt bridge creates a barrier to contamination which minimises the dilution of the internal standard cell solution. A replaceable salt bridge holds a large volume of buffer to protect the reference electrode from harsh process conditions. The salt bridge threads onto the end of the sensor if replacement is needed.

Encapsulated construction protects the sensor’s built-in preamp from moisture and humidity. A preamp in the pHD analogue sensor produces a strong enough signal, enabling the sensor to be located up to 1000 m from the controller.

Includes the capability to remotely monitor sensors on any browser-enabled device and present diagnostics on the overall health of the measurements (on predictive diagnostics-enabled sensors), as well as upcoming and required maintenance - reducing user risk and downtime.

Technical Data

Measuring range Accuracy Sensitivity Repeatability

Temperature accuracy

Drift Operating

temperature range

0 - 14 pH

± 0.02 pH

± 0.01 pH

± 0.05 pH

± 0.5 °C

0.03 pH per 24 hours, non-cumulative Digital Sensor: -5 - 70 °C

Analogue Sensor with Digital Gateway: 5 - 105 °C Immersion Sensor: 0 - 50 °C

3 m per second, maximum

Max. 2 bar overpressure

Flow rate

Sensor pressure / Temperature limit

Temperature sensor

NTC 300 Ω thermistor for automatic temperature compensation and analyser temperature

readout

100 m, maximum

Transmission distance

Cable length Wetted materials Storage conditions Calibration method Immersion depth Communication Weight

10 m PUR (polyurethane) 4-conductor with one shield, rated to 105 °C Stainless steel, Ryton, glass, titanium, FKM/FPM o-ring

4 - 70 °C, 0 - 95 % relative humidity (non-condensing)

Two point automatic, one point automatic, two point manual, one point manual Submersible to 107 m/1050 kPa

Modbus

0.870 kg

PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Scope of delivery

Sensor comes with a 10 m cable and manual

Electrode Type

(Select at least 1 option)

[ ] General purpose glass electrode for most applications

[ ] Glass electrode for resistance to HF, low pH applications

Sensor Body Material (Select at least 1 option)

[ ] PEEK body (for chemical compatibility with most process solutions) [ ] Ryton body (for less aggressive solutions)

[ ] Stainless steel body (for immersion applications).

Sensor Mounting Options (Select at least 1 option)

[ ] Convertible [ ] Insertion

[ ] Immersion [ ] Sanitary

Accessories

[ ] (delete what is not needed)

[ ] Universal controller

[ ] Buffer solution, pH 7.00 (NIST), colour-coded yellow, 500 mL [ ] Buffer solution, pH 4.01 (NIST), colour-coded red, 500 mL

[ ] Buffer solution, pH 10.01 (NIST), colour-coded blue, 500 mL

[ ] Standard cell solution, concentrated pH 7.0 buffer (equi-transferrant), 500 mL [ ] PEEK salt bridge, Kynar (PVDF) outer-junction

[ ] Digital extension cable, 7.7 m [ ] Digital extension cable, 15 m [ ] Digital extension cable, 31 m

[ ] Pole mount assembly for 1” NPT sensors [ ] CPVC Union Mount for pHD sensor

[ ] Float mount assembly for 1” NPT sensors

[ ] CPVC flow-through mount assembly hardware for pHD differential sensors [ ] Insertion mounting, digital, 1 1/2” CPVC

[ ] Self-contained air blast cleaning systems (115 V AC) for pHD

Services

[ ] (delete what is not needed)

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:

Hach

Product:

pHD sc Digital Online Process pH Sensor