PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Stationary Automatic Sampler

Fixed site sampler in stainless steel housing with thermostatic control for automatic sample extraction according to pressure-vacuum sampling principle. Mains operation 230 V/50 Hz or 115 V/60 Hz.

Technical Data

Housing material 	Double-walled stainless steel (material 1.4301/SS304) / PS / PC (GF10) with 40 mm insulation. Housing separated in sample compartment and control compartment, each with lockable door. Upper door with Plexiglas window. Protective top made of Styrosun which can be opened for connection and maintenance works.

Optional: 1.4571/SS316Ti; SS304 Epoxy-coated; SS316Ti Epoxy-coated

Cooling system	Self-contained, controlled cooling / heating with 4 settings, no-frost. independent of the programmable controller, temperature in sample compartment: 4 °C (adjustable from 0.0 -

9.9 °C)

Controller	Microprocessor control, Sleep-Mode (<5 mA), power supply 8-16 V, foil keyboard (with keys

0 - 9, ESC, ENT, cursor), graphical display (128 x 64 Pixel), backlit

Data storage	3000 entries, non-volatile data memory; storage of sampling and malfunction data like sample extractions, bottle changes, messages, external signals

Programs	12 freely programmable user programs, with function to link programs

Program start options	Immediately; at a certain time; by an external signal

Program stop options	End of sampling program after one program run; continuous operation or x- runs, and a certain date / time

Pause mode 	Interruption of program run at any time Overfilling protection	Adjustable from 1 - 999 samples/bottle Sample interval	1 min to 99 h 59 min in steps of 1 minute

1 to 9999 pulses/sample

Manual sampling	Possible at any time without interrupting the current program run

Program protection	Up to 5 years after voltage loss

Interface 	Mini-USB

Communication	Optional: Modbus, connection via Profibus DP, PC via Software

Further communication options on request.

Languages user interface

Multi-language, selectable

Inputs	1 x analogue: 0/4-20 mA

2 x digital (flow, event), Impulse length 60 ms, switching level 7-24 V, max. working

resistance 500 Ohm, max. length of signal cable 30 m

Status messages	Standard: 1 collective malfunction message as relay output

Option: expandable up to 5 relay outputs - freely programmable

Sampling principle	Standard Vacuum System 20 - 350 mL Option: Vacuum System 20 - 500 mL

Option: Vacuum CTVV flow-proportional 5 - 350 mL

Option: Bypass System 20 - 250 mL

Option: glass dosing vessel instead of plastic

Sample volume accuracy

Standard Vacuum System: < 2.5 % or ± 3 mL

Suction height	Max. 8 m (at 1013 hPa and stagnant medium), optional higher

Sample transport velocity

>0.5 m/s at suction height up to at least 8 m (at 1013 hPa); pump performance can be adjusted electronically

Suction hose material 	PVC

Suction hose length	7.5 m; max. hose length 30 m

Suction hose inner diameter

12 mm

Sampling mode 	Time-related, flow-proportional, constant time/ variable volume (CTVV), event-related,

manual sample extraction

Sample container	Plastic: 1 x 25 L; 1 x 50 L; 2 x 10 L; 2 x 22 L; 4 x 10 L; 4 x 14 L; 4 x 20 L; 4 x 25 L; 12 x 2.9

L; 24 x 1,0 L; 24 x 2.9 L

Glass: 12 x 2.0 L; 24 x 0.9 L; 24 x 2.0 L

Dimensions	1290 (1890*) mm x 690 mm x 645 mm (H x W x D)

* with opened top

Weight	Approx. 100 kg with composite container, higher weight when using multi bottle option and/

or glass bottles

Power requirements

(Voltage)

115/230 V AC

Power supply	Approx. 350 VA (with cooling)

Ambient temperature 	-20 - 43 °C

Sample temperature 	0 - 40 °C

Certifications	CE, Sampling according to ISO 5667-2/3-10

Material 	PC, PVC, Silicone, PS, PE, EPDM (optional: metering vessel glass Duran50, sinker weight

SS304)

Scope of delivery

[ ] (delete what is not needed)

[ ] 7.5 m PVC suction hose, stainless steel strainer

[ ] Basic Controller (includes 0/4-20 mA, digital impulses, event Signal input port, collective malfunction relay and mini-USB port, 350 mL Plastic dosing vessel)

[ ] Basic Controller (includes 0/4-20 mA, digital impulses, event Signal input port, collective malfunction relay and mini-USB port, 350 mL Glass dosing vessel)

[ ] Basic Controller (includes 0/4-20 mA, digital impulses, event Signal input port, collective malfunction relay and mini-USB port, 350 mL Plastic dosing vessel for CTVV sampling mode)

[ ] Basic Controller (includes 0/4-20 mA, digital impulses, event Signal input port, collective malfunction relay and mini-USB port, 500 mL Glass dosing vessel)

[ ] Basic Controller (includes 0/4-20 mA, digital impulses, event Signal input port, collective malfunction relay and mini-USB port, 250 mL Bypass Glass dosing vessel for pressurized line sampling)

[ ] Controller (includes 0/4-20 mA, digital impulses, event Signal input port, I/O extention with 4 relay outputs and mini-USB port, sample compartment interior lightning, 350 mL Plastic dosing vessel)

[ ] Controller (includes 0/4-20 mA, digital impulses, event Signal input port, I/O extention with 4 relay outputs and mini-USB port, sample compartment interior lightning, 350 mL Glass dosing vessel)

[ ] Controller (includes 0/4-20 mA, digital impulses, event Signal input port, I/O extention with 4 relay outputs and mini-USB port, sample compartment interior lightning, 350 mL Plastic dosing vessel for CTVV sampling mode)

[ ] Controller (includes 0/4-20 mA, digital impulses, event Signal input port, I/O extention with 4 relay outputs and mini-USB port, sample compartment interior lightning, 500 mL Glass dosing vessel)

[ ] Controller (includes 0/4-20 mA, digital impulses, event Signal input port, I/O extention with 4 relay outputs and

mini-USB port, sample compartment interior lightning, 250 mL Bypass Glass dosing vessel for pressurized line sampling)

Bottle Sets

[ ] Composite Plastic, 1 x 25 L [ ] Composite Plastic, 1 x 50 L [ ] Discrete Plastic, 2 x 10 L

[ ] Discrete Plastic, 2 x 22 L [ ] Discrete Plastic, 4 x 10 L [ ] Discrete Plastic, 4 x 14 L [ ] Discrete Plastic, 4 x 20 L [ ] Discrete Plastic, 4 x 25 L

[ ] Discrete Plastic, 12 x 2.9 L [ ] Discrete Plastic, 24 x 1 L

[ ] Discrete Plastic, 24 x 2.9 L [ ] Discrete Glass, 12 x 2 L

[ ] Discrete Glass, 24 x 0.9 L [ ] Discrete Glass, 24 x 2 L

Intake Tubing

[ ] 7.5 m suction hose (12.7 mm ID), PVC, counterweight stainless steel

Accessories

[ ] Overvoltage protection for Signal input

[ ] Overvoltage protection for mains

[ ] Built-in option: capacitive level sensor instead of conductive

[ ] Built-in option: handles for Bühler 401x to lift sampler

[ ] Built-in option: Profibus DP communication

[ ] Built-in option: Modbus communication

[ ] Built-in option: Webserver communication via RJ45 Interface

[ ] Built-in option: non-return valve (slanted valve) [ ] Built-in option: backup battery with charger

[ ] Built-in option: main switch

[ ] Suction hose protection heater band (length 5 m) [ ] Castors (rollers)

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Hach

Product:	Bühler 4011 Stationary Automatic Sampler