PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Conductivity Sensor

Contacting conductivity sensor with integrated temperature probe for use in a variety of applications over a broad range of conductivities. These applications include pure and ultrapure water, moderately conductive solutions, and waste water and drinking water. The sensor is designed to work with a controller for data collection and operation, and is factory- tested to determine its absolute cell constant and temperature constant for accuracy.

Technical Data

Contacting conductivity sensor, low range, model 1

Application	Pure and ultra pure water Material	Head/Body: Polysulfone/SS Cell constant k	0.01 cm-1

Measuring range              0.01 - 200 µS/cm Operating temperature    125 °C maximum Pressure	10 bar maximum

Contacting conductivity sensor, low range, model 2

Application	Pure and ultra pure water Material	Head/Body: SS/SS316L Cell constant k	0.01 cm-1

Measuring range              0.01 - 200 µS/cm Operating temperature    150 °C maximum Pressure	25 bar maximum

Contacting conductivity sensor, medium range, model 3

Application                       Moderately conductive solutions

Material                             Head/Body: Polysulfone/SS

Cell constant k                 0.1 cm-1

Measuring range              0.1 - 2000 µS/cm Operating temperature    125 °C maximum Pressure	10 bar maximum

Contacting conductivity sensor, medium range, model 4

Application                       Moderately conductive solutions

Material                             Head/Body: SS/SS316L

Cell constant k                 0.1 cm-1

Measuring range              0.1 - 2000 µS/cm Operating temperature    150 °C maximum Pressure	25 bar maximum

Contacting conductivity sensor, high range, model 5

Application	Wastewater & Drinking water Material	Head/Body: Polysulfone/Graphite Cell constant k	1.00 cm-1

Measuring range	1 - 20000 µS/cm Operating temperature	125 °C maximum Pressure	10 bar maximum

Contacting conductivity sensor, high range, model 6

Application	Wastewater & Drinking water

Material	Head/Body: SS/Graphite

Cell constant k	1.00 cm-1

Measuring range	1 - 20000 µS/cm Operating temperature	150 °C maximum Pressure	25 bar maximum

Contacting conductivity sensor (in general, for all ranges and models) Cable connection	Fiber-glass-reinforced polyester plug head Accuracy	< 2 %

Response time	< 30 s

Process connection	3/4“ NPT thread

Mounting	Bypass or Inline

Scope of delivery

Sensor and user manual

Electrode type

[ ] (select at least 1 option)

[ ] Cell constant: 0.01 cm-1, range: 0.01-200 µS/cm, SS electrodes, 10 bar/125 °C (model 1)

[ ] Cell constant: 0.01 cm-1, range: 0.01-200 µS/cm, SS 316L electrodes, 25 bar/150 °C (model 2) [ ] Cell constant: 0.1 cm-1, range: 0.01-2,000 µS/cm, SS electrodes, 10 bar/125 °C (model 3)

[ ] Cell constant: 0.1 cm-1, range: 0.01-2,000 µS/cm, SS 316L electrodes, 25 bar/150 °C (model 4)

[ ] Cell constant: 1.0 cm-1, range: 0.01-20,000 µS/cm, Graphite electrodes, 10 bar/125 °C (model 5) [ ] Cell constant: 1.0 cm-1, range: 0.01-20,000 µS/cm, Graphite electrodes, 25 bar/150 °C (model 6)

Accessories

[ ] (delete what is not needed)

[ ] Female connector 6+T with connection drawing

[ ] 5 m cable and IP65 connector for 2 electrode conductivity sensor

[ ] 10 m cable and IP65 connector for 2 electrode conductivity sensor [ ] 20 m cable and IP65 connector for 2 electrode conductivity sensor [ ] 30 m cable and IP65 connector for 2 electrode conductivity sensor [ ] PVC flow chamber with 3 x ¾ FNPT bores

[ ] Stainless steel flow chamber with 1 X ¾ FNPT bore + 2 x ¼ FNPT bores

[ ] Kit for 8394 1½ inch clamp probe with EPDM gasket, clamp and 316L SS flow chamber

[ ] Kit for 8394 2 inch clamp probe with EPDM gasket, clamp and 316L flow chamber

Services

[ ] (delete what is not needed)

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:	Polymetron

Product:	8310, 8311, 8312, 8315, 8316, 8317 Conductivity Sensors