Single or multi-channel Online Total Organic Carbon Analyser

Analyser for measuring Total Organic Carbon, Total Carbon, Total Inorganic Carbon, Biological Oxygen Demand and/ or Chemical Oxygen Demand via correlation, and Volatile Organic Carbon via calculation. The analyser is designed for measurement of these parameters in water. The Total Organic Carbon measurement method is infrared measurement of Carbon Dioxide after oxidation. The oxidation method is Two Stage Advanced Oxidation which utilises the Hydroxyl radical and a manganese catalyst.

The analyser is capable of measuring Total Organic Carbon every 6.5 minutes. Up to 6 process streams. The sample sequencer is internal to the analyser.

The analyser is self-cleaning. The analyser uses the internal sample pump to bring the sample into the reactor. The analyser reactor is self-cleaning. The analyser has internal sensors and utilises self-diagnostics to monitor analyser status and produces alerts when maintenance is required.

Technical Data

Parameter

Direct measurement of Total Organic Carbon, Total Inorganic Carbon, Total Carbon Chemical Oxygen Demand, Biological Oxygen Demand via correlation, Volatile Organic Carbon via calculation

Infrared measurement of CO2 after oxidation

Unique Two-Stage Advanced Oxidation Process using Hydroxyl Radicals 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

Automatic or Manual Range Selection 1, 2, 4 and 6 streams

±3 % of reading or ±0.45 mg/L C, whichever is greater in 0 - 250 mg/L C range

±3 % of reading or ±6 mg/L C, whichever is greater in 0 - 2000 mg/L C range

±3 % of reading or ±30 mg/L C, whichever is greater in 0 - 20000 mg/L C range From 6.5 minutes, depending on application

Up to 30%

Measurement method Oxidation method Measuring range Range selection

Multi-Stream

Repeatability

Cycle time

Permissible Chloride range

Communication: digital

Enclosure waterproof rating

Sample inlet temperature Ambient temperature

Modbus RTU, Modbus TCP/IP & Profibus

IP44; optional IP54 with air purge

2 - 60 °C

5 - 40 °C

Air conditioning and heating options are available. 5 - 85 % (non-condensing)

Up to 2 mm, soft particulates

Previous 9999 analysis data on screen in the microcontroller memory and storage of data archive for the lifetime of the analyser in the SD/MMC card.

Previous 99 fault data on screen in the microcontroller memory and storage of fault data archive for the lifetime of the analyser in the SD/MMC card.

High contrast 40 character x 16 line backlit LCD with CFL backlight 115 V AC / 230 V AC

Humidity Particle size

Data storage

Display

Power requirements (Voltage)

Power requirements (Hz) Service interval Dimensions (H x W x D) Weight

50/60 Hz

6 months service intervals 1250 mm x 750 mm x 320 mm

90 - 120 kg

Enclosure weight may change depending on system optional features.

PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.

Scope of delivery

[ ] (delete what is not needed)

Analytical Instrument

[ ] 1 channel Online TOC Analyser, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

[ ] 1 channel Online TOC Analyser, IP54 Purge Ready, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C [ ] 2 channel Online TOC Analyser, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

[ ] 2 channel Online TOC Analyser, IP54 Purge Ready, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C [ ] 4 channel Online TOC Analyser, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

[ ] 4 channel Online TOC Analyser, IP54 Purge Ready, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C [ ] 6 channel Online TOC Analyser, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

[ ] 6 channel Online TOC Analyser, IP54 Purge Ready, 0 - 250 mg/L C, 0 - 2000 mg/L C, 0 - 20000 mg/L C

Accessories and Options

[ ] VOC mode used for measurement of Volatile Organic Carbon in the sample [ ] Remote activation of Manual/Calibration port

[ ] Profibus

[ ] Modbus RTU

[ ] Modbus TCP/IP

[ ] Sample sensor to detect presence of the sample and evaluate the quantity of air bubbles in the sample

[ ] Sample overflow chamber

[ ] Vacuum sampler [ ] Venturi Sampler

[ ] Flow-through sampling chamber with sand trap [ ] Sampling chamber with sand trap

[ ] Compressor

[ ] Filter pack for removal of water, dirt, and oil from compressed air (if using instrument air supply)

Services

[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of

performance of the instrument.

[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

Brand:

Hach

Product:

B7000i Dairy Online TOC Analyser